## instructions
Introduction
This workbook is designed to help users calculate radiation exposure after shielding, or determine the shielding thickness required.
It includes
(a) more than 20 radioisotopes,
(b) 3 shielding materials ( lead, concrete, iron),
(c) both 2π and 4π geometries,
(d) 2 calculation types : (exposure after shielding or thickness of shielding required)
(e) Calculations with or without build-up factors.
Only the yellow cells are interactive. All other cells are protected to prevent accidental modification.
Step 1 : select inputs
click the yellow cells and select from the dropdown menu :
Radioisotope
Shielding naterial
Geometry
Calculation type
with/without build up factor
The background or font color will change depending on your selection
Step 2 : enter numeric values
in the calculation box below , enter the numeric value of
(a) initial exposure
(b) target exposure or shielding thickness, depending on the calculation type selected
The calculator performs all computations automatically
Example 1 : | Example 2 :
Calculate exposure after a given shielding thickness : | Calculate shielding thickness required :
Tc-99m initial exposure =200 μSv/hr, lead shielding thickness =0.05 cm, 2π geometry without build up factor | F-18 exposure =250 μSv/hr, target exposure = 0.5 μSv/hr, 4π geometry with build up factor
Calculated exposure = 49.99 μSv/hr | Calculated concrete thickness required = 50.68 cm

## Disclaimer
Disclaimer
The contents of this spreadsheet are provided solely for educational and reference purposes. They are intended to assist understanding and preliminary exploration only.
These tools do not replace site-specific engineering design, institutional standard operating procedures, professional consultancy, or regulatory review and approval.
As with all calculation tools, users should independently verify all results with appropriately qualified professionals and ensure compliance with applicable institutional protocols, standards, and regulatory requirements.
Regulatory requirements, dose limits, and shielding standards vary by country and jurisdiction.Users are therefore responsible for consulting relevant national regulations and competent authorities before applying any results.
The tools provided on this website are not intended to be used as the sole basis for decision-making.Users are solely responsible for verifying all inputs, assumptions, units, methodologies, and outputs prior to any practical application.
The author assumes no responsibility or liability for decisions made based on the use of these tools, nor for any errors arising from incorrect data entry, inappropriate assumptions, misinterpretation, or misuse of the contents.

## Sheet1
|  | Radiation shielding Calculator |  | copyright © Sidney Yu |  | radioisotope list | shielding material | calculation type list
|  |  |  | please refer to Disclaimer page before downloading or using this worksheet |  |  | List
select | radioisotope: | Fluorine -18 | [5 empty] | exposure after shielding
[6 empty] | Carbon-11 | lead | thickness of shielding required
select | shielding material : | concrete |  |  |  | Cesium -137 | concrete
[6 empty] | Cobalt -57 | iron
select | geometry: | 4π |  | When build up factor is not considered , based on the shielding equation, |  | Cobalt-60
|  |  |  | shielding geometry does not affect shielding calculation |  | Copper -64
select | calculation type: | thickness of shielding required |  |  |  | Copper- 67 | geometry list | thickness of shielding required | target exposure = | µSv/hr | thickness required= | cm
Fluorine -18
select | calculation with /without build up factor : |  |  | When lead is the shielding material AND the radioisotope energy |  | Gallium -68 | 2π | exposure after shielding | shielding thickness = | cm | final exposure = | µSv / hr
|  | with |  | is less than 0.3 MeV, there is no build up factor |  | Iodine -123 | 4π | thickness of shielding required | target exposure = | µSv/hr | thickness required= | cm
Iodine - 124
Fluorine -18 | concrete |  | 4π |  |  | Iodine-131
[6 empty] | Iridium -192 | with/without build up factor
| calculate | thickness of shielding required | with | build up factor |  | Lutetium -177
[6 empty] | Nitrogen -13 | with
enter | initial exposure = | 250 | µSv/hr |  |  | Radium -223 | without
Rhenium-186
enter | target exposure = | 0.5 | µSv/hr |  |  | Rhenium -188
Rubidium -82
| thickness required= | 50.6837 | cm | with | build up factor | Technetium-99m
Zirconium-89
Yttrium-90

## Sheet2 (hidden)
convert to codes
radioisotope: | Fluorine -18 |  |  | shielding material code : |  | 2000 | calculation type code : | 20
shielding material : | concrete |  |  | lead | 1000 |  | exposure after shielding | 10
|  |  |  | concrete | 2000 |  | thickness of shielding required | 20
geometry: | 4π |  |  | iron | 3000
calculation type: | thickness of shielding required
|  |  |  | geomerty code : |  | 200 | with/without build up factor code: | 1
calculation with /without build up factor :
| with |  |  | 2π | 100 |  | with | 1
|  |  |  | 4π | 200 |  | without | 2
|  |  |  | combined | code = | 2221
|  | linear attenuation coefficent column |  |  |  | 12 | build up table search addition code::
|  | lead | 11 |  |  |  | concrete4π | 30
|  | concrete | 12
|  | iron | 13 |  |  |  | lead2π | 0
[7 empty] | lead4π | 10
[7 empty] | concrete2π | 20
[7 empty] | concrete4π | 30
[7 empty] | iron2π | 40
[7 empty] | iron4π | 50

## Sheet3 (hidden)
search for linear attenuation coefficient and first colun in build up factor table
radioisotope: | Fluorine -18
shielding material : | concrete
shielding material coulmn : |  | 12
| µ = | 0.2032 | cm-1
first column in build up factor table = |  | 6
| MFP = | 4.9213 | cm | MFP = mean free path
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13
Name | symbol | half life |  | primary decay mode | Energy | highest gamma | corresponding | half value layer (lead) | specific gamma constant | Linear attenuation coefficient Lead | Linear attenuation coefficient concrete | Linear attenuation coefficient Iron
[5 empty] | (MeV) |  | column | (cm) | (mSv/hr)/MBq | cm-1 | cm-1 | cm-1
Actinium -225 | Ac-225 | 9.97 | day | alpha α | α = 5.8 - 8.4 |  |  | 0.00147 | 5.172e-05 | 472.6
Carbon-11 | C-11 | 20.5 | min | positron β+ | 0.511 | 0.511 | 6 | 0.41 | 0.000191 | 1.69 | 0.2032 | 0.6625 | 0.6625
Cesium -137 | Cs-137 | 30.1 | year | gamma γ | 0.661 | 0.661 | 7 | 0.57 | 0.000108 | 1.27 | 0.1811 | 0.5787 | 0.5787
Cobalt -57 | Co-57 | 270.9 | day | gamma γ | 0.122, 0.136 | 0.136 | 2 | 0.018 | 4.09e-05 | 39.81 | 0.3478 | 1.878 | 1.878
Cobalt-60 | Co-60 | 5.27 | year | gamma γ | 1.17, 1.33 | 1.33 | 9 | 1.02 | 0.0003697 | 0.679 | 0.1293 | 0.4109 | 0.4109
Copper -64 | Cu-64 | 12.7 | hour | positron β+ Beta β- | γ = 0.511 βmax = 0.578 βave = 0.190 | 0.511 | 6 | 0.6 | 3.51e-05 | 1.614 | 0.2032 | 0.6625 | 0.6625
Copper- 67 | Cu-67 | 61.8 | hr | gamma γ beta β- | γ = 0.185 β1 max = 0.377 ave = 0.121 β2 max = 0.408 ave = 0.154 β3 max = 0.562 ave= 0.189 | 0.185 | 3 | < 0.1 | 2.63e-05 | 13.6929 | 0.3037 | 1.2347 | 1.2347
Fluorine -18 | F-18 | 1.8 | hour | positron β+ | 0.511 | 0.511 | 6 | 0.41 | 0.000185 | 1.695 | 0.2032 | 0.6625 | 0.6625
Gallium -68 | Ga-68 | 1.1 | hour | positron β+ | 0.511 | 0.511 | 6 | 0.43 | 0.000176 | 1.608 | 0.2032 | 0.6625 | 0.6625
Iodine -123 | I-123 | 13.1 | hour | gamma γ | 0.159 | 0.159 | 3 | 0.036 | 7.48e-05 | 19.42 | 0.3246 | 1.5469 | 1.5469
Iodine -124 | I-124 | 4.2 | day | positron β+ beta β- |  | 1.69 | 10 | 0.8 | 0.0002033 | 0.969 | 0.1145 | 0.3759 | 0.3759
Iodine-131 | I-131 | 8 | day | gamma γ beta β- | γ = 0.364 βmax = 0.61 βave = 0.192 | 0.364 | 5 | 0.29 | 7.64e-05 | 2.41 | 0.2332 | 0.7896 | 0.7896
Iridium -192 | Ir-192 | 74 | day | gamma γ | 0.296, 0.612, 1.4 | 1.4 | 9 | 0.44 | 0.00016 | 2.334 | 0.126 | 0.401 | 0.401
Lutetium -177 | Lu-177 | 6.7 | day | gamma γ beta β- | γ = 0.208 βmax = 0.498 βave = 0.149 | 0.208 | 3 | 0.06 | 7.64e-06 | 11.55 | 0.2949 | 1.1496 | 1.1496
Nitrogen -13 | N-13 | 10 | min | positron β+ | 0.511 | 0.511 | 6 | 0.6 | 0.000191 | 1.695 | 0.2032 | 0.6625 | 0.6625
Radium -223 | Ra-223 | 11.4 | day | alpha α | α = 5.6 -6.7 γ = 0.084, 0.269 | 0.269 | 4 | 0.064 | 8.78e-05 | 10.79 | 0.2629 | 0.9476 | 0.9476
Rhenium-186 | Re-186 | 3.8 | day | gamma γ beta β- | γ = 0.102 βmax = 1.077 βave = 0.362 | 0.102 | 2 | 0.03 | 4.91e-06 | 28.58 | 0.3997 | 2.9268 | 2.9268
Rhenium -188 | Re-188 | 17 | hour | gamma γ beta β- | γ = 0.155 βmax = 2.12 βave = 0.784 | 0.155 | 3 | 0.52 | 1.09e-05 | 1.325 | 0.3303 | 1.5465 | 1.5465
Rubidium -82 | Rb-82 | 1.26 | min | positron β+ | 0.511 | 0.511 | 6 | 0.35 | 0.000208 | 1.571 | 0.2032 | 0.6625 | 0.6625
Technetium-99m | Tc-99m | 6.02 | hour | gamma γ | 0.14 | 0.14 | 2 | 0.025 | 3.32e-05 | 27.73 | 0.3434 | 1.8053 | 1.8053
Zirconium-89 | Zr-89 | 3.3 | day | positron β+ | 0.511 | 0.511 | 6 | 0.66 | 0.000266 | 1.045 | 0.2032 | 0.6625 | 0.6625
Terbium-161 | Tb-161 | 6.95 | day | gamma γ beta β- | γ = 0.049, 0.0746 βmax = 0.593 βave = 0.154 | 0.0746 | 2 | not avaliable | not avaliable | not available
Yttrium-90 | Y-90 | 64.2 | hour | beta β- | βmax = 2.28 βave = 0.93 |  |  |  | 2.44e-08 | 776.6 |  |  | 776.6
0.6625
Name
Actinium -225
Carbon-11 | 0.6625
Cesium -137 | 0.5787
Cobalt -57 | 1.878
Cobalt-60 | 0.4109
Copper -64 | 0.6625
Copper- 67 | 1.2347
Fluorine -18 | 0.6625
Gallium -68 | 0.6625
Iodine -123 | 1.5469
Iodine -124 | 0.3759
Iodine-131 | 0.7896
Iridium -192 | 0.401
Lutetium -177 | 1.1496
Nitrogen -13 | 0.6625
Radium -223 | 0.9476
Rhenium-186 | 2.9268
Rhenium -188 | 1.5465
Rubidium -82 | 0.6625
Technetium-99m | 1.8053
Zirconium-89 | 0.6625
Terbium-161
Yttrium-90

## Sheet4 (hidden)
search for build up factor
|  |  |  | calculated #MFP | rounded #MFP |  |  | lead | 2π | [9 empty] | lead | 4π | [8 empty] | Concrete | 2π | [8 empty] | Concrete | 4π | [8 empty] | Iron | 2π | [8 empty] | Iron | 4π
| radioisotope: | Fluorine -18 |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60
|  |  |  | 0.1 | 0.5 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2
| shielding material : | concrete |  | 0.2 | 0.5 | 0.5 | 1 | 1 | 1 | 1.11 | 1.14 | 1.16 | 1.2 | 1.21 | 1.25 |  | 1 | 1 | 1 | 1.19 | 1.24 | 1.29 | 1.35 | 1.42 | 1.48 |  | 2.15 | 2.05 | 1.93 | 1.84 | 1.76 | 1.71 | 1.83 | 1.56 | 1.43 |  | 2.76 | 2.68 | 2.52 | 2.4 | 2.3 | 2.22 | 2.11 | 2.02 | 1.82 |  | 1.26 | 1.53 | 1.6 | 1.61 | 1.59 | 1.57 | 1.53 | 1.49 | 1.4 |  | 1.4 | 1.88 | 2.03 | 2.07 | 2.05 | 2.02 | 1.96 | 1.91 | 1.77
|  |  |  | 0.3 | 0.5 | 1 | 1 | 1 | 1 | 1.17 | 1.22 | 1.27 | 1.33 | 1.37 | 1.44 |  | 1 | 1 | 1 | 1.29 | 1.37 | 1.44 | 1.55 | 1.66 | 1.77 |  | 2.74 | 2.72 | 2.56 | 2.43 | 2.32 | 2.24 | 2.11 | 2.01 | 1.79 |  | 4.17 | 4.21 | 3.9 | 3.67 | 3.46 | 3.3 | 3.07 | 2.89 | 2.44 |  | 1.36 | 1.83 | 1.99 | 2.03 | 2.02 | 2 | 1.93 | 1.88 | 1.73 |  | 1.61 | 2.48 | 2.8 | 2.89 | 2.86 | 2.82 | 2.72 | 2.62 | 2.32
| geometry: | 4π |  | 0.4 | 0.5 | 2 | 1 | 1 | 1 | 1.26 | 1.35 | 1.43 | 1.55 | 1.64 | 1.79 |  | 1 | 1 | 1 | 1.44 | 1.57 | 1.7 | 1.89 | 2.08 | 2.3 |  | 3.85 | 4.16 | 3.92 | 3.73 | 3.54 | 3.38 | 3.14 | 2.96 | 2.52 |  | 7.32 | 8.05 | 7.4 | 6.81 | 6.32 | 5.92 | 5.32 | 4.86 | 3.74 |  | 1.51 | 2.36 | 2.73 | 2.87 | 2.88 | 2.85 | 2.77 | 2.68 | 2.39 |  | 1.97 | 3.65 | 4.43 | 4.69 | 4.68 | 4.61 | 4.39 | 4.17 | 3.46
|  |  |  | 0.5 | 0.5 | 3 | 1 | 1 | 1 | 1.33 | 1.46 | 1.57 | 1.75 | 1.87 | 2.13 |  | 1 | 1 | 1 | 1.56 | 1.75 | 1.93 | 2.2 | 2.47 | 2.83 |  | 4.99 | 5.78 | 5.48 | 5.21 | 4.95 | 4.66 | 4.31 | 4.02 | 3.28 |  | 10.9 | 13 | 12 | 10.9 | 9.94 | 9.17 | 8.03 | 7.18 | 5.13 |  | 1.64 | 2.86 | 3.48 | 3.75 | 3.82 | 3.76 | 3.68 | 3.54 | 3.09 |  | 2.28 | 4.87 | 6.24 | 6.75 | 6.76 | 6.67 | 6.31 | 5.93 | 4.69
| calculation type: | thickness of shielding required |  | 0.6 | 1 | 4 | 1 | 1 | 1 | 1.39 | 1.55 | 1.68 | 1.92 | 2.09 | 2.47 |  | 1 | 1 | 1 | 1.67 | 1.91 | 2.14 | 2.49 | 2.83 | 3.35 |  | 6.16 | 7.6 | 7.28 | 6.9 | 8.55 | 6.09 | 5.61 | 5.17 | 4.08 |  | 15 | 19.1 | 17.6 | 15.8 | 14.3 | 13.1 | 11.2 | 9.81 | 6.59 |  | 1.74 | 3.34 | 4.25 | 4.68 | 4.82 | 4.74 | 4.66 | 4.46 | 3.82 |  | 2.56 | 6.11 | 8.22 | 9.06 | 9.12 | 9 | 8.47 | 7.87 | 5.98
|  |  |  | 0.7 | 1 | 5 | 1 | 1 | 1 | 1.44 | 1.63 | 1.79 | 2.08 | 2.29 | 2.8 |  | 1 | 1 | 1 | 1.78 | 2.07 | 2.34 | 2.76 | 3.18 | 3.87 |  | 7.4 | 9.68 | 9.33 | 8.83 | 8.38 | 7.71 | 7.06 | 6.43 | 4.92 |  | 19.6 | 26.6 | 24.6 | 21.8 | 19.5 | 17.6 | 14.8 | 12.7 | 8.13 |  | 1.84 | 3.8 | 5.03 | 5.66 | 5.91 | 5.87 | 5.72 | 5.44 | 4.58 |  | 2.82 | 7.37 | 10.3 | 11.6 | 11.8 | 11.6 | 10.9 | 10 | 7.36
|  |  |  | 0.8 | 1 | 6 | 1 | 1 | 1 | 1.49 | 1.7 | 1.89 | 2.23 | 2.48 | 3.14 |  | 1 | 1 | 1 | 1.87 | 2.21 | 2.53 | 3.03 | 3.53 | 4.42 |  | 8.76 | 12.1 | 11.7 | 11.1 | 10.5 | 9.54 | 8.69 | 7.82 | 5.82 |  | 24.8 | 35.7 | 33 | 29 | 25.7 | 23 | 18.9 | 16.1 | 9.83 |  | 1.91 | 4.22 | 5.78 | 6.63 | 7.11 | 7.07 | 6.89 | 6.51 | 5.4 |  | 3.05 | 8.58 | 12.5 | 14.2 | 14.8 | 14.6 | 13.6 | 12.4 | 8.86
| type code = | 20 |  | 0.9 | 1 | 7 | 1 | 1 | 1 | 1.53 | 1.77 | 1.98 | 2.37 | 2.67 | 3.48 |  | 1 | 1 | 1 | 1.96 | 2.35 | 2.72 | 3.29 | 3.86 | 4.96 |  | 10.2 | 14.7 | 14.4 | 13.6 | 12.8 | 11.5 | 10.5 | 9.3 | 6.75 |  | 30.5 | 46.3 | 42.8 | 37.3 | 32.7 | 29 | 23.5 | 19.7 | 11.6 |  | 1.99 | 4.64 | 6.57 | 7.68 | 8.38 | 8.35 | 8.14 | 7.63 | 6.24 |  | 3.28 | 9.86 | 14.8 | 17 | 18 | 17.8 | 16.5 | 15 | 10.4
|  |  |  | 1 | 1 | 8 | 1 | 1 | 1 | 1.57 | 1.83 | 2.06 | 2.51 | 2.84 | 3.81 |  | 1 | 1 | 1 | 2.05 | 2.49 | 2.89 | 3.54 | 4.18 | 5.51 |  | 11.7 | 17.7 | 17.3 | 16.3 | 15.4 | 13.7 | 12.4 | 10.9 | 7.71 |  | 36.7 | 58.6 | 54.1 | 46.7 | 40.6 | 35.7 | 28.5 | 23.6 | 13.4 |  | 2.06 | 5.06 | 7.38 | 8.77 | 9.74 | 9.71 | 9.47 | 8.82 | 7.12 |  | 3.5 | 11.2 | 17.2 | 20.1 | 21.5 | 21.3 | 19.7 | 17.7 | 12
| MFP = | 4.9213 |  | 1.1 | 1 | 10 | 1 | 1 | 1 | 1.64 | 1.94 | 2.21 | 2.76 | 3.18 | 4.49 |  | 1 | 1 | 1 | 2.2 | 2.73 | 3.23 | 4.02 | 4.8 | 6.62 |  | 14.9 | 24.5 | 24.2 | 22.7 | 21.4 | 18.7 | 16.7 | 14.3 | 9.7 |  | 50.5 | 88.4 | 81.6 | 69.2 | 59 | 51.1 | 39.7 | 32 | 17.1 |  | 2.19 | 5.9 | 9.07 | 11.1 | 12.7 | 12.6 | 12.4 | 11.3 | 8.95 |  | 3.92 | 13.9 | 22.5 | 26.9 | 29.3 | 29.1 | 26.6 | 23.6 | 15.4
|  |  |  | 1.2 | 1 | 15 | 1 | 1 | 1 | 1.78 | 2.17 | 2.54 | 3.33 | 3.93 | 6.18 |  | 1 | 1 | 1 | 2.53 | 3.3 | 3.99 | 5.12 | 6.24 | 9.45 |  | 24.3 | 48 | 47.7 | 44.2 | 41 | 34.2 | 30 | 24.3 | 15.1 |  | 93.8 | 199 | 182 | 148 | 121 | 101 | 74.4 | 57 | 27.1 |  | 2.48 | 8.01 | 13.6 | 17.9 | 21.5 | 21.5 | 21 | 18.5 | 13.9 |  | 4.91 | 21 | 37.6 | 47.1 | 53.2 | 53 | 47.5 | 40.6 | 24.7
| shielding thickess = | not applicable | cm | 1.3 | 1 | 20 | 1 | 1 | 1 | 1.89 | 2.37 | 2.82 | 3.84 | 4.61 | 7.89 |  | 1 | 1 | 1 | 2.8 | 3.82 | 4.68 | 6.12 | 7.58 | 12.3 |  | 35.8 | 82 | 81 | 74.1 | 67.5 | 54.2 | 46.9 | 36.1 | 21 |  | 149 | 371 | 333 | 262 | 205 | 168 | 118 | 86.6 | 38 |  | 2.74 | 10.1 | 18.6 | 25.7 | 32.3 | 32.4 | 31.4 | 26.8 | 19.4 |  | 5.83 | 28.6 | 55.1 | 71.7 | 82.7 | 82.7 | 72.8 | 60.5 | 34.9
|  |  |  | 1.4 | 1 | 25 | 1 | 1 | 1 | 1.99 | 2.56 | 3.07 | 4.3 | 5.23 | 9.62 |  | 1 | 1 | 1 | 3.04 | 4.31 | 5.34 | 7.07 | 8.85 | 15.3 |  | 49.1 | 128 | 125 | 113 | 101 | 78.7 | 67 | 49.4 | 27.2 |  | 217 | 615 | 541 | 412 | 311 | 250 | 170 | 120 | 49.4 |  | 2.97 | 12.3 | 23.9 | 34.7 | 44.9 | 45.2 | 43.5 | 35.8 | 25.4 |  | 6.71 | 36.5 | 74.6 | 100 | 117 | 118 | 102 | 82.6 | 45.7
| calculated # of MFP = | #VALUE! |  | 1.5 | 2 | 30 | 1 | 1 | 1 | 2.08 | 2.73 | 3.31 | 4.72 | 5.83 | 11.3 |  | 1 | 1 | 1 | 3.26 | 4.79 | 5.96 | 7.97 | 10.1 | 18.2 |  | 64.3 | 189 | 181 | 161 | 142 | 108 | 90.4 | 63.9 | 33.8 |  | 295 | 942 | 810 | 600 | 437 | 347 | 229 | 157 | 61.3 |  | 3.18 | 14.5 | 29.7 | 44.7 | 59.1 | 60 | 57.3 | 45.6 | 31.6 |  | 7.56 | 44.6 | 95.6 | 132 | 156 | 157 | 135 | 106 | 57.1
|  |  |  | 1.6 | 2 | 35 | 1 | 1 | 1 | 2.16 | 2.89 | 3.53 | 5.12 | 6.39 | 13.1 |  | 1 | 1 | 1 | 3.45 | 5.25 | 6.56 | 8.83 | 11.2 | 21.2 |  | 81 | 264 | 249 | 220 | 189 | 141 | 117 | 79.3 | 40.6 |  | 384 | 1360 | 1140 | 826 | 582 | 457 | 294 | 196 | 73.5 |  | 3.38 | 16.7 | 35.7 | 55.6 | 74.8 | 76.7 | 72.6 | 55.9 | 38.1 |  | 8.39 | 53 | 118 | 167 | 199 | 201 | 171 | 132 | 68.9
| round up # of MFP = | #VALUE! |  | 1.7 | 2 | 40 | 1 | 1 | 1 | 2.24 | 3.05 | 3.74 | 5.49 | 6.93 | 14.8 |  | 1 | 1 | 1 | 3.63 | 5.71 | 7.15 | 9.66 | 12.4 | 24.2 |  | 99.2 | 357 | 330 | 288 | 242 | 177 | 146 | 95.6 | 47.5 |  | 492 | 1870 | 1540 | 1090 | 745 | 579 | 365 | 237 | 85.9 |  | 3.58 | 19 | 42 | 67.4 | 91.8 | 95.1 | 89.2 | 66.6 | 44.9 |  | 9.22 | 61.4 | 141 | 205 | 245 | 248 | 210 | 159 | 80.9
|  |  |  | 1.8 | 2
| first column in build |  |  | 1.9 | 2
| up factor table = | 6 |  | 2 | 2
|  |  |  | 2.1 | 2
| column additor = | 30 |  | 2.2 | 2
|  |  |  | 2.3 | 2
| column to search |  |  | 2.4 | 2
| in build up factor table = | 36 |  | 2.5 | 3
|  |  |  | 2.6 | 3
| raw build up factor = | #VALUE! |  | 2.7 | 3
|  |  |  | 2.8 | 3
| adjusted build up factor= | #VALUE! |  | 2.9 | 3
|  |  |  | 3 | 3
|  |  |  | 3.1 | 3
NOTE: | current calculation is find thickness, this page is only to find build up factor, refer to sheet6 for calculation |  |  | 3.2 | 3
|  |  |  | 3.3 | 3
|  |  |  | 3.4 | 3
|  |  |  | 3.5 | 4
|  |  |  | 3.6 | 4
|  |  |  | 3.7 | 4
|  |  |  | 3.8 | 4
|  |  |  | 3.9 | 4
|  |  |  | 4 | 4
|  |  |  | 4.1 | 4
|  |  |  | 4.2 | 4
|  |  |  | 4.3 | 4
|  |  |  | 4.4 | 4
|  |  |  | 4.5 | 5
|  |  |  | 4.6 | 5
|  |  |  | 4.7 | 5
|  |  |  | 4.8 | 5
|  |  |  | 4.9 | 5
|  |  |  | 5 | 5
|  |  |  | 5.1 | 5
|  |  |  | 5.2 | 5
|  |  |  | 5.3 | 5
|  |  |  | 5.4 | 5
|  |  |  | 5.5 | 6
|  |  |  | 5.6 | 6
|  |  |  | 5.7 | 6
|  |  |  | 5.8 | 6
|  |  |  | 5.9 | 6
|  |  |  | 6 | 6
|  |  |  | 6.1 | 6
|  |  |  | 6.2 | 6
|  |  |  | 6.3 | 6
|  |  |  | 6.4 | 6
|  |  |  | 6.5 | 7
|  |  |  | 6.6 | 7
|  |  |  | 6.7 | 7
|  |  |  | 6.8 | 7
|  |  |  | 6.9 | 7
|  |  |  | 7 | 7
|  |  |  | 7.1 | 7
|  |  |  | 7.2 | 7
|  |  |  | 7.3 | 7
|  |  |  | 7.4 | 7
|  |  |  | 7.5 | 8
|  |  |  | 7.6 | 8
|  |  |  | 7.7 | 8
|  |  |  | 7.8 | 8
|  |  |  | 7.9 | 8
|  |  |  | 8 | 8
|  |  |  | 8.1 | 8
|  |  |  | 8.2 | 8
|  |  |  | 8.3 | 8
|  |  |  | 8.4 | 8
|  |  |  | 8.5 | 8
|  |  |  | 8.6 | 8
|  |  |  | 8.7 | 8
|  |  |  | 8.8 | 8
|  |  |  | 8.9 | 8
|  |  |  | 9 | 10
|  |  |  | 9.1 | 10
|  |  |  | 9.2 | 10
|  |  |  | 9.3 | 10
|  |  |  | 9.4 | 10
|  |  |  | 9.5 | 10
|  |  |  | 9.6 | 10
|  |  |  | 9.7 | 10
|  |  |  | 9.8 | 10
|  |  |  | 9.9 | 10
|  |  |  | 10 | 10
|  |  |  | 10.1 | 10
|  |  |  | 10.2 | 10
|  |  |  | 10.3 | 10
|  |  |  | 10.4 | 10
|  |  |  | 10.5 | 10
|  |  |  | 10.6 | 10
|  |  |  | 10.7 | 10
|  |  |  | 10.8 | 10
|  |  |  | 10.9 | 10
|  |  |  | 11 | 10
|  |  |  | 11.1 | 10
|  |  |  | 11.2 | 10
|  |  |  | 11.3 | 10
|  |  |  | 11.4 | 10
|  |  |  | 11.5 | 10
|  |  |  | 11.6 | 10
|  |  |  | 11.7 | 10
|  |  |  | 11.8 | 10
|  |  |  | 11.9 | 10
|  |  |  | 12 | 10
|  |  |  | 12.1 | 10
|  |  |  | 12.2 | 10
|  |  |  | 12.3 | 10
|  |  |  | 12.4 | 10
|  |  |  | 12.5 | 15
|  |  |  | 12.6 | 15
|  |  |  | 12.7 | 15
|  |  |  | 12.8 | 15
|  |  |  | 12.9 | 15
|  |  |  | 13 | 15
|  |  |  | 13.1 | 15
|  |  |  | 13.2 | 15
|  |  |  | 13.3 | 15
|  |  |  | 13.4 | 15
|  |  |  | 13.5 | 15
|  |  |  | 13.6 | 15
|  |  |  | 13.7 | 15
|  |  |  | 13.8 | 15
|  |  |  | 13.9 | 15
|  |  |  | 14 | 15
|  |  |  | 14.1 | 15
|  |  |  | 14.2 | 15
|  |  |  | 14.3 | 15
|  |  |  | 14.4 | 15
|  |  |  | 14.5 | 15
|  |  |  | 14.6 | 15
|  |  |  | 14.7 | 15
|  |  |  | 14.8 | 15
|  |  |  | 14.9 | 15
|  |  |  | 15 | 15
|  |  |  | 15.1 | 15
|  |  |  | 15.2 | 15
|  |  |  | 15.3 | 15
|  |  |  | 15.4 | 15
|  |  |  | 15.5 | 15
|  |  |  | 15.6 | 15
|  |  |  | 15.7 | 15
|  |  |  | 15.8 | 15
|  |  |  | 15.9 | 15
|  |  |  | 16 | 15
|  |  |  | 16.1 | 15
|  |  |  | 16.2 | 15
|  |  |  | 16.3 | 15
|  |  |  | 16.4 | 15
|  |  |  | 16.5 | 15
|  |  |  | 16.6 | 15
|  |  |  | 16.7 | 15
|  |  |  | 16.8 | 15
|  |  |  | 16.9 | 15
|  |  |  | 17 | 15
|  |  |  | 17.1 | 15
|  |  |  | 17.2 | 15
|  |  |  | 17.3 | 15
|  |  |  | 17.4 | 15
|  |  |  | 17.5 | 20
|  |  |  | 17.6 | 20
|  |  |  | 17.7 | 20
|  |  |  | 17.8 | 20
|  |  |  | 17.9 | 20
|  |  |  | 18 | 20
|  |  |  | 18.1 | 20
|  |  |  | 18.2 | 20
|  |  |  | 18.3 | 20
|  |  |  | 18.4 | 20
|  |  |  | 18.5 | 20
|  |  |  | 18.6 | 20
|  |  |  | 18.7 | 20
|  |  |  | 18.8 | 20
|  |  |  | 18.9 | 20
|  |  |  | 19 | 20
|  |  |  | 19.1 | 20
|  |  |  | 19.2 | 20
|  |  |  | 19.3 | 20
|  |  |  | 19.4 | 20
|  |  |  | 19.5 | 20
|  |  |  | 19.6 | 20
|  |  |  | 19.7 | 20
|  |  |  | 19.8 | 20
|  |  |  | 19.9 | 20
|  |  |  | 20 | 20
|  |  |  | 20.1 | 20
|  |  |  | 20.2 | 20
|  |  |  | 20.3 | 20
|  |  |  | 20.4 | 20
|  |  |  | 20.5 | 20
|  |  |  | 20.6 | 20
|  |  |  | 20.7 | 20
|  |  |  | 20.8 | 20
|  |  |  | 20.9 | 20
|  |  |  | 21 | 20
|  |  |  | 21.1 | 20
|  |  |  | 21.2 | 20
|  |  |  | 21.3 | 20
|  |  |  | 21.4 | 20
|  |  |  | 21.5 | 20
|  |  |  | 21.6 | 20
|  |  |  | 21.7 | 20
|  |  |  | 21.8 | 20
|  |  |  | 21.9 | 20
|  |  |  | 22 | 20
|  |  |  | 22.1 | 20
|  |  |  | 22.2 | 20
|  |  |  | 22.3 | 20
|  |  |  | 22.4 | 20
|  |  |  | 22.5 | 25
|  |  |  | 22.6 | 25
|  |  |  | 22.7 | 25
|  |  |  | 22.8 | 25
|  |  |  | 22.9 | 25
|  |  |  | 23 | 25
|  |  |  | 23.1 | 25
|  |  |  | 23.2 | 25
|  |  |  | 23.3 | 25
|  |  |  | 23.4 | 25
|  |  |  | 23.5 | 25
|  |  |  | 23.6 | 25
|  |  |  | 23.7 | 25
|  |  |  | 23.8 | 25
|  |  |  | 23.9 | 25
|  |  |  | 24 | 25
|  |  |  | 24.1 | 25
|  |  |  | 24.2 | 25
|  |  |  | 24.3 | 25
|  |  |  | 24.4 | 25
|  |  |  | 24.5 | 25
|  |  |  | 24.6 | 25
|  |  |  | 24.7 | 25
|  |  |  | 24.8 | 25
|  |  |  | 24.9 | 25
|  |  |  | 25 | 25
|  |  |  | 25.1 | 25
|  |  |  | 25.2 | 25
|  |  |  | 25.3 | 25
|  |  |  | 25.4 | 25
|  |  |  | 25.5 | 25
|  |  |  | 25.6 | 25
|  |  |  | 25.7 | 25
|  |  |  | 25.8 | 25
|  |  |  | 25.9 | 25
|  |  |  | 26 | 25
|  |  |  | 26.1 | 25
|  |  |  | 26.2 | 25
|  |  |  | 26.3 | 25
|  |  |  | 26.4 | 25
|  |  |  | 26.5 | 25
|  |  |  | 26.6 | 25
|  |  |  | 26.7 | 25
|  |  |  | 26.8 | 25
|  |  |  | 26.9 | 25
|  |  |  | 27 | 25
|  |  |  | 27.1 | 25
|  |  |  | 27.2 | 25
|  |  |  | 27.3 | 25
|  |  |  | 27.4 | 25
|  |  |  | 27.5 | 30
|  |  |  | 27.6 | 30
|  |  |  | 27.7 | 30
|  |  |  | 27.8 | 30
|  |  |  | 27.9 | 30
|  |  |  | 28 | 30
|  |  |  | 28.1 | 30
|  |  |  | 28.2 | 30
|  |  |  | 28.3 | 30
|  |  |  | 28.4 | 30
|  |  |  | 28.5 | 30
|  |  |  | 28.6 | 30
|  |  |  | 28.7 | 30
|  |  |  | 28.8 | 30
|  |  |  | 28.9 | 30
|  |  |  | 29 | 30
|  |  |  | 29.1 | 30
|  |  |  | 29.2 | 30
|  |  |  | 29.3 | 30
|  |  |  | 29.4 | 30
|  |  |  | 29.5 | 30
|  |  |  | 29.6 | 30
|  |  |  | 29.7 | 30
|  |  |  | 29.8 | 30
|  |  |  | 29.9 | 30
|  |  |  | 30 | 30
|  |  |  | 30.1 | 30
|  |  |  | 30.2 | 30
|  |  |  | 30.3 | 30
|  |  |  | 30.4 | 30
|  |  |  | 30.5 | 30
|  |  |  | 30.6 | 30
|  |  |  | 30.7 | 30
|  |  |  | 30.8 | 30
|  |  |  | 30.9 | 30
|  |  |  | 31 | 30
|  |  |  | 31.1 | 30
|  |  |  | 31.2 | 30
|  |  |  | 31.3 | 30
|  |  |  | 31.4 | 30
|  |  |  | 31.5 | 30
|  |  |  | 31.6 | 30
|  |  |  | 31.7 | 30
|  |  |  | 31.8 | 30
|  |  |  | 31.9 | 30
|  |  |  | 32 | 30
|  |  |  | 32.1 | 30
|  |  |  | 32.2 | 30
|  |  |  | 32.3 | 30
|  |  |  | 32.4 | 30
|  |  |  | 32.5 | 35
|  |  |  | 32.6 | 35
|  |  |  | 32.7 | 35
|  |  |  | 32.8 | 35
|  |  |  | 32.9 | 35
|  |  |  | 33 | 35
|  |  |  | 33.1 | 35
|  |  |  | 33.2 | 35
|  |  |  | 33.3 | 35
|  |  |  | 33.4 | 35
|  |  |  | 33.5 | 35
|  |  |  | 33.6 | 35
|  |  |  | 33.7 | 35
|  |  |  | 33.8 | 35
|  |  |  | 33.9 | 35
|  |  |  | 34 | 35
|  |  |  | 34.1 | 35
|  |  |  | 34.2 | 35
|  |  |  | 34.3 | 35
|  |  |  | 34.4 | 35
|  |  |  | 34.5 | 35
|  |  |  | 34.6 | 35
|  |  |  | 34.7 | 35
|  |  |  | 34.8 | 35
|  |  |  | 34.9 | 35
|  |  |  | 35 | 35
|  |  |  | 35.1 | 35
|  |  |  | 35.2 | 35
|  |  |  | 35.3 | 35
|  |  |  | 35.4 | 35
|  |  |  | 35.5 | 35
|  |  |  | 35.6 | 35
|  |  |  | 35.7 | 35
|  |  |  | 35.8 | 35
|  |  |  | 35.9 | 35
|  |  |  | 36 | 35
|  |  |  | 36.1 | 35
|  |  |  | 36.2 | 35
|  |  |  | 36.3 | 35
|  |  |  | 36.4 | 35
|  |  |  | 36.5 | 35
|  |  |  | 36.6 | 35
|  |  |  | 36.7 | 35
|  |  |  | 36.8 | 35
|  |  |  | 36.9 | 35
|  |  |  | 37 | 35
|  |  |  | 37.1 | 35
|  |  |  | 37.2 | 35
|  |  |  | 37.3 | 35
|  |  |  | 37.4 | 35
|  |  |  | 37.5 | 40
|  |  |  | 37.6 | 40
|  |  |  | 37.7 | 40
|  |  |  | 37.8 | 40
|  |  |  | 37.9 | 40
|  |  |  | 38 | 40
|  |  |  | 38.1 | 40
|  |  |  | 38.2 | 40
|  |  |  | 38.3 | 40
|  |  |  | 38.4 | 40
|  |  |  | 38.5 | 40
|  |  |  | 38.6 | 40
|  |  |  | 38.7 | 40
|  |  |  | 38.8 | 40
|  |  |  | 38.9 | 40
|  |  |  | 39 | 40
|  |  |  | 39.1 | 40
|  |  |  | 39.2 | 40
|  |  |  | 39.3 | 40
|  |  |  | 39.4 | 40
|  |  |  | 39.5 | 40
|  |  |  | 39.6 | 40
|  |  |  | 39.7 | 40
|  |  |  | 39.8 | 40
|  |  |  | 39.9 | 40
|  |  |  | 40 | 40
|  |  |  | 40.1 | 40
|  |  |  | 40.2 | 40
|  |  |  | 40.3 | 40
|  |  |  | 40.4 | 40
|  |  |  | 40.5 | 40
|  |  |  | 40.6 | 40
|  |  |  | 40.7 | 40
|  |  |  | 40.8 | 40
|  |  |  | 40.9 | 40
|  |  |  | 41 | 40
|  |  |  | 41.1 | 40
|  |  |  | 41.2 | 40
|  |  |  | 41.3 | 40
|  |  |  | 41.4 | 40
|  |  |  | 41.5 | 40
|  |  |  | 41.6 | 40
|  |  |  | 41.7 | 40
|  |  |  | 41.8 | 40
|  |  |  | 41.9 | 40
|  |  |  | 42 | 40
|  |  |  | 42.1 | 40
|  |  |  | 42.2 | 40
|  |  |  | 42.3 | 40
|  |  |  | 42.4 | 40

## Sheet5 (hidden)
calculation : exposure after shielding with /without building up factor
| radioisotope: | Fluorine -18 | combined code = | 2221
| shielding material : | concrete | initial exposure = | 250
| geometry: | 4π | shielding thickness= | 0.5
| calculation with /without build up factor : |  | W or W/O build up factor code = | 1
with
|  |  | build up factor = | #VALUE!
|  |  | linear attenuation coef= | 0.2032
|  |  | final expoure = | #VALUE!
NOTE: | current calculation to find thickness, this page not in operation, refer to sheet 6

## Sheet6 (hidden)
calculation : thickness of shielding required
select | radioisotope: | Fluorine -18
select | shielding material : | concrete |  |  | combined code = |  | 2221 | [5 empty] | calculated #MFP | rounded #MFP |  |  | lead | 2π | [9 empty] | lead | 4π | [8 empty] | Concrete | 2π | [8 empty] | Concrete | 4π | [8 empty] | Iron | 2π | [8 empty] | Iron | 4π
[15 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60
select | geometry: | 4π |  | W or W/O build up factor code= |  |  | 1 | [5 empty] | 0.1 | 0.5 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2 | # of MFP | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.8 | 1 | 2
[13 empty] | 0.2 | 0.5 | 0.5 | 1 | 1 | 1 | 1.11 | 1.14 | 1.16 | 1.2 | 1.21 | 1.25 |  | 1 | 1 | 1 | 1.19 | 1.24 | 1.29 | 1.35 | 1.42 | 1.48 |  | 2.15 | 2.05 | 1.93 | 1.84 | 1.76 | 1.71 | 1.83 | 1.56 | 1.43 |  | 2.76 | 2.68 | 2.52 | 2.4 | 2.3 | 2.22 | 2.11 | 2.02 | 1.82 |  | 1.26 | 1.53 | 1.6 | 1.61 | 1.59 | 1.57 | 1.53 | 1.49 | 1.4 |  | 1.4 | 1.88 | 2.03 | 2.07 | 2.05 | 2.02 | 1.96 | 1.91 | 1.77
select | calculation with /without build up factor : |  |  | prelimnary thickness W/O build up factor= |  |  |  | 30.5837 | cm |  |  |  | 0.3 | 0.5 | 1 | 1 | 1 | 1 | 1.17 | 1.22 | 1.27 | 1.33 | 1.37 | 1.44 |  | 1 | 1 | 1 | 1.29 | 1.37 | 1.44 | 1.55 | 1.66 | 1.77 |  | 2.74 | 2.72 | 2.56 | 2.43 | 2.32 | 2.24 | 2.11 | 2.01 | 1.79 |  | 4.17 | 4.21 | 3.9 | 3.67 | 3.46 | 3.3 | 3.07 | 2.89 | 2.44 |  | 1.36 | 1.83 | 1.99 | 2.03 | 2.02 | 2 | 1.93 | 1.88 | 1.73 |  | 1.61 | 2.48 | 2.8 | 2.89 | 2.86 | 2.82 | 2.72 | 2.62 | 2.32
|  | with |  | thickness with build up factor = |  |  | 50.6837 | cm |  |  |  |  | 0.4 | 0.5 | 2 | 1 | 1 | 1 | 1.26 | 1.35 | 1.43 | 1.55 | 1.64 | 1.79 |  | 1 | 1 | 1 | 1.44 | 1.57 | 1.7 | 1.89 | 2.08 | 2.3 |  | 3.85 | 4.16 | 3.92 | 3.73 | 3.54 | 3.38 | 3.14 | 2.96 | 2.52 |  | 7.32 | 8.05 | 7.4 | 6.81 | 6.32 | 5.92 | 5.32 | 4.86 | 3.74 |  | 1.51 | 2.36 | 2.73 | 2.87 | 2.88 | 2.85 | 2.77 | 2.68 | 2.39 |  | 1.97 | 3.65 | 4.43 | 4.69 | 4.68 | 4.61 | 4.39 | 4.17 | 3.46
|  |  |  | Final result -thicness required with or W/O build up factor = | [5 empty] | 50.6837 |  |  | 0.5 | 0.5 | 3 | 1 | 1 | 1 | 1.33 | 1.46 | 1.57 | 1.75 | 1.87 | 2.13 |  | 1 | 1 | 1 | 1.56 | 1.75 | 1.93 | 2.2 | 2.47 | 2.83 |  | 4.99 | 5.78 | 5.48 | 5.21 | 4.95 | 4.66 | 4.31 | 4.02 | 3.28 |  | 10.9 | 13 | 12 | 10.9 | 9.94 | 9.17 | 8.03 | 7.18 | 5.13 |  | 1.64 | 2.86 | 3.48 | 3.75 | 3.82 | 3.76 | 3.68 | 3.54 | 3.09 |  | 2.28 | 4.87 | 6.24 | 6.75 | 6.76 | 6.67 | 6.31 | 5.93 | 4.69
[13 empty] | 0.6 | 1 | 4 | 1 | 1 | 1 | 1.39 | 1.55 | 1.68 | 1.92 | 2.09 | 2.47 |  | 1 | 1 | 1 | 1.67 | 1.91 | 2.14 | 2.49 | 2.83 | 3.35 |  | 6.16 | 7.6 | 7.28 | 6.9 | 8.55 | 6.09 | 5.61 | 5.17 | 4.08 |  | 15 | 19.1 | 17.6 | 15.8 | 14.3 | 13.1 | 11.2 | 9.81 | 6.59 |  | 1.74 | 3.34 | 4.25 | 4.68 | 4.82 | 4.74 | 4.66 | 4.46 | 3.82 |  | 2.56 | 6.11 | 8.22 | 9.06 | 9.12 | 9 | 8.47 | 7.87 | 5.98
| initial exposure = | 250 |  | estimation of thickness W build up factor: | [8 empty] | 0.7 | 1 | 5 | 1 | 1 | 1 | 1.44 | 1.63 | 1.79 | 2.08 | 2.29 | 2.8 |  | 1 | 1 | 1 | 1.78 | 2.07 | 2.34 | 2.76 | 3.18 | 3.87 |  | 7.4 | 9.68 | 9.33 | 8.83 | 8.38 | 7.71 | 7.06 | 6.43 | 4.92 |  | 19.6 | 26.6 | 24.6 | 21.8 | 19.5 | 17.6 | 14.8 | 12.7 | 8.13 |  | 1.84 | 3.8 | 5.03 | 5.66 | 5.91 | 5.87 | 5.72 | 5.44 | 4.58 |  | 2.82 | 7.37 | 10.3 | 11.6 | 11.8 | 11.6 | 10.9 | 10 | 7.36
[13 empty] | 0.8 | 1 | 6 | 1 | 1 | 1 | 1.49 | 1.7 | 1.89 | 2.23 | 2.48 | 3.14 |  | 1 | 1 | 1 | 1.87 | 2.21 | 2.53 | 3.03 | 3.53 | 4.42 |  | 8.76 | 12.1 | 11.7 | 11.1 | 10.5 | 9.54 | 8.69 | 7.82 | 5.82 |  | 24.8 | 35.7 | 33 | 29 | 25.7 | 23 | 18.9 | 16.1 | 9.83 |  | 1.91 | 4.22 | 5.78 | 6.63 | 7.11 | 7.07 | 6.89 | 6.51 | 5.4 |  | 3.05 | 8.58 | 12.5 | 14.2 | 14.8 | 14.6 | 13.6 | 12.4 | 8.86
| target exposure = | 0.5 |  | thickness increment= |  | 0.1 | cm | [5 empty] | 0.9 | 1 | 7 | 1 | 1 | 1 | 1.53 | 1.77 | 1.98 | 2.37 | 2.67 | 3.48 |  | 1 | 1 | 1 | 1.96 | 2.35 | 2.72 | 3.29 | 3.86 | 4.96 |  | 10.2 | 14.7 | 14.4 | 13.6 | 12.8 | 11.5 | 10.5 | 9.3 | 6.75 |  | 30.5 | 46.3 | 42.8 | 37.3 | 32.7 | 29 | 23.5 | 19.7 | 11.6 |  | 1.99 | 4.64 | 6.57 | 7.68 | 8.38 | 8.35 | 8.14 | 7.63 | 6.24 |  | 3.28 | 9.86 | 14.8 | 17 | 18 | 17.8 | 16.5 | 15 | 10.4
[13 empty] | 1 | 1 | 8 | 1 | 1 | 1 | 1.57 | 1.83 | 2.06 | 2.51 | 2.84 | 3.81 |  | 1 | 1 | 1 | 2.05 | 2.49 | 2.89 | 3.54 | 4.18 | 5.51 |  | 11.7 | 17.7 | 17.3 | 16.3 | 15.4 | 13.7 | 12.4 | 10.9 | 7.71 |  | 36.7 | 58.6 | 54.1 | 46.7 | 40.6 | 35.7 | 28.5 | 23.6 | 13.4 |  | 2.06 | 5.06 | 7.38 | 8.77 | 9.74 | 9.71 | 9.47 | 8.82 | 7.12 |  | 3.5 | 11.2 | 17.2 | 20.1 | 21.5 | 21.3 | 19.7 | 17.7 | 12
| linear attenuation coef= | 0.2032 | cm-1 | target difference between exposure and target exp= |  |  |  |  | 0 |  |  |  | 1.1 | 1 | 10 | 1 | 1 | 1 | 1.64 | 1.94 | 2.21 | 2.76 | 3.18 | 4.49 |  | 1 | 1 | 1 | 2.2 | 2.73 | 3.23 | 4.02 | 4.8 | 6.62 |  | 14.9 | 24.5 | 24.2 | 22.7 | 21.4 | 18.7 | 16.7 | 14.3 | 9.7 |  | 50.5 | 88.4 | 81.6 | 69.2 | 59 | 51.1 | 39.7 | 32 | 17.1 |  | 2.19 | 5.9 | 9.07 | 11.1 | 12.7 | 12.6 | 12.4 | 11.3 | 8.95 |  | 3.92 | 13.9 | 22.5 | 26.9 | 29.3 | 29.1 | 26.6 | 23.6 | 15.4
[13 empty] | 1.2 | 1 | 15 | 1 | 1 | 1 | 1.78 | 2.17 | 2.54 | 3.33 | 3.93 | 6.18 |  | 1 | 1 | 1 | 2.53 | 3.3 | 3.99 | 5.12 | 6.24 | 9.45 |  | 24.3 | 48 | 47.7 | 44.2 | 41 | 34.2 | 30 | 24.3 | 15.1 |  | 93.8 | 199 | 182 | 148 | 121 | 101 | 74.4 | 57 | 27.1 |  | 2.48 | 8.01 | 13.6 | 17.9 | 21.5 | 21.5 | 21 | 18.5 | 13.9 |  | 4.91 | 21 | 37.6 | 47.1 | 53.2 | 53 | 47.5 | 40.6 | 24.7
| MFP = | 4.9213 |  | thickness | cal.#MFP | round#MDP | build up factr | adj B | exposu | diff |  | thickness | 1.3 | 1 | 20 | 1 | 1 | 1 | 1.89 | 2.37 | 2.82 | 3.84 | 4.61 | 7.89 |  | 1 | 1 | 1 | 2.8 | 3.82 | 4.68 | 6.12 | 7.58 | 12.3 |  | 35.8 | 82 | 81 | 74.1 | 67.5 | 54.2 | 46.9 | 36.1 | 21 |  | 149 | 371 | 333 | 262 | 205 | 168 | 118 | 86.6 | 38 |  | 2.74 | 10.1 | 18.6 | 25.7 | 32.3 | 32.4 | 31.4 | 26.8 | 19.4 |  | 5.83 | 28.6 | 55.1 | 71.7 | 82.7 | 82.7 | 72.8 | 60.5 | 34.9
|  |  |  | 30.6837 | 6.2349 | 6 | 25.7 | 25.7 | 12.5915 | 12.0915 | -- | 50.6837 | 1.4 | 1 | 25 | 1 | 1 | 1 | 1.99 | 2.56 | 3.07 | 4.3 | 5.23 | 9.62 |  | 1 | 1 | 1 | 3.04 | 4.31 | 5.34 | 7.07 | 8.85 | 15.3 |  | 49.1 | 128 | 125 | 113 | 101 | 78.7 | 67 | 49.4 | 27.2 |  | 217 | 615 | 541 | 412 | 311 | 250 | 170 | 120 | 49.4 |  | 2.97 | 12.3 | 23.9 | 34.7 | 44.9 | 45.2 | 43.5 | 35.8 | 25.4 |  | 6.71 | 36.5 | 74.6 | 100 | 117 | 118 | 102 | 82.6 | 45.7
| first build up factor table column = | 6 |  | 30.7837 | 6.2552 | 6 | 25.7 | 25.7 | 12.3382 | 11.8382 | -- |  | 1.5 | 2 | 30 | 1 | 1 | 1 | 2.08 | 2.73 | 3.31 | 4.72 | 5.83 | 11.3 |  | 1 | 1 | 1 | 3.26 | 4.79 | 5.96 | 7.97 | 10.1 | 18.2 |  | 64.3 | 189 | 181 | 161 | 142 | 108 | 90.4 | 63.9 | 33.8 |  | 295 | 942 | 810 | 600 | 437 | 347 | 229 | 157 | 61.3 |  | 3.18 | 14.5 | 29.7 | 44.7 | 59.1 | 60 | 57.3 | 45.6 | 31.6 |  | 7.56 | 44.6 | 95.6 | 132 | 156 | 157 | 135 | 106 | 57.1
|  |  |  | 30.8837 | 6.2756 | 6 | 25.7 | 25.7 | 12.0901 | 11.5901 | -- |  | 1.6 | 2 | 35 | 1 | 1 | 1 | 2.16 | 2.89 | 3.53 | 5.12 | 6.39 | 13.1 |  | 1 | 1 | 1 | 3.45 | 5.25 | 6.56 | 8.83 | 11.2 | 21.2 |  | 81 | 264 | 249 | 220 | 189 | 141 | 117 | 79.3 | 40.6 |  | 384 | 1360 | 1140 | 826 | 582 | 457 | 294 | 196 | 73.5 |  | 3.38 | 16.7 | 35.7 | 55.6 | 74.8 | 76.7 | 72.6 | 55.9 | 38.1 |  | 8.39 | 53 | 118 | 167 | 199 | 201 | 171 | 132 | 68.9
| build up factor table additor= | 30 |  | 30.9837 | 6.2959 | 6 | 25.7 | 25.7 | 11.8469 | 11.3469 | -- |  | 1.7 | 2 | 40 | 1 | 1 | 1 | 2.24 | 3.05 | 3.74 | 5.49 | 6.93 | 14.8 |  | 1 | 1 | 1 | 3.63 | 5.71 | 7.15 | 9.66 | 12.4 | 24.2 |  | 99.2 | 357 | 330 | 288 | 242 | 177 | 146 | 95.6 | 47.5 |  | 492 | 1870 | 1540 | 1090 | 745 | 579 | 365 | 237 | 85.9 |  | 3.58 | 19 | 42 | 67.4 | 91.8 | 95.1 | 89.2 | 66.6 | 44.9 |  | 9.22 | 61.4 | 141 | 205 | 245 | 248 | 210 | 159 | 80.9
|  |  |  | 31.0837 | 6.3162 | 6 | 25.7 | 25.7 | 11.6086 | 11.1086 | -- |  | 1.8 | 2
| build up factor search column= | 36 |  | 31.1837 | 6.3365 | 6 | 25.7 | 25.7 | 11.3751 | 10.8751 | -- |  | 1.9 | 2
|  |  |  | 31.2837 | 6.3568 | 6 | 25.7 | 25.7 | 11.1463 | 10.6463 | -- |  | 2 | 2
|  |  |  | 31.3837 | 6.3772 | 6 | 25.7 | 25.7 | 10.9221 | 10.4221 | -- |  | 2.1 | 2
|  |  |  | 31.4837 | 6.3975 | 6 | 25.7 | 25.7 | 10.7024 | 10.2024 | -- |  | 2.2 | 2
|  |  |  | 31.5837 | 6.4178 | 6 | 25.7 | 25.7 | 10.4871 | 9.9871 | -- |  | 2.3 | 2
|  |  |  | 31.6837 | 6.4381 | 6 | 25.7 | 25.7 | 10.2761 | 9.7761 | -- |  | 2.4 | 2
|  |  |  | 31.7837 | 6.4584 | 6 | 25.7 | 25.7 | 10.0694 | 9.5694 | -- |  | 2.5 | 3
|  |  |  | 31.8837 | 6.4788 | 6 | 25.7 | 25.7 | 9.8669 | 9.3669 | -- |  | 2.6 | 3
|  |  |  | 31.9837 | 6.4991 | 6 | 25.7 | 25.7 | 9.6684 | 9.1684 | -- |  | 2.7 | 3
|  |  |  | 32.0837 | 6.5194 | 7 | 32.7 | 32.7 | 12.0544 | 11.5544 | -- |  | 2.8 | 3
|  |  |  | 32.1837 | 6.5397 | 7 | 32.7 | 32.7 | 11.8119 | 11.3119 | -- |  | 2.9 | 3
|  |  |  | 32.2837 | 6.56 | 7 | 32.7 | 32.7 | 11.5743 | 11.0743 | -- |  | 3 | 3
|  |  |  | 32.3837 | 6.5804 | 7 | 32.7 | 32.7 | 11.3415 | 10.8415 | -- |  | 3.1 | 3
|  |  |  | 32.4837 | 6.6007 | 7 | 32.7 | 32.7 | 11.1134 | 10.6134 | -- |  | 3.2 | 3
|  |  |  | 32.5837 | 6.621 | 7 | 32.7 | 32.7 | 10.8898 | 10.3898 | -- |  | 3.3 | 3
|  |  |  | 32.6837 | 6.6413 | 7 | 32.7 | 32.7 | 10.6708 | 10.1708 | -- |  | 3.4 | 3
|  |  |  | 32.7837 | 6.6616 | 7 | 32.7 | 32.7 | 10.4561 | 9.9561 | -- |  | 3.5 | 4
|  |  |  | 32.8837 | 6.682 | 7 | 32.7 | 32.7 | 10.2458 | 9.7458 | -- |  | 3.6 | 4
|  |  |  | 32.9837 | 6.7023 | 7 | 32.7 | 32.7 | 10.0397 | 9.5397 | -- |  | 3.7 | 4
|  |  |  | 33.0837 | 6.7226 | 7 | 32.7 | 32.7 | 9.8378 | 9.3378 | -- |  | 3.8 | 4
|  |  |  | 33.1837 | 6.7429 | 7 | 32.7 | 32.7 | 9.6399 | 9.1399 | -- |  | 3.9 | 4
|  |  |  | 33.2837 | 6.7632 | 7 | 32.7 | 32.7 | 9.446 | 8.946 | -- |  | 4 | 4
|  |  |  | 33.3837 | 6.7836 | 7 | 32.7 | 32.7 | 9.256 | 8.756 | -- |  | 4.1 | 4
|  |  |  | 33.4837 | 6.8039 | 7 | 32.7 | 32.7 | 9.0698 | 8.5698 | -- |  | 4.2 | 4
|  |  |  | 33.5837 | 6.8242 | 7 | 32.7 | 32.7 | 8.8873 | 8.3873 | -- |  | 4.3 | 4
|  |  |  | 33.6837 | 6.8445 | 7 | 32.7 | 32.7 | 8.7086 | 8.2086 | -- |  | 4.4 | 4
|  |  |  | 33.7837 | 6.8648 | 7 | 32.7 | 32.7 | 8.5334 | 8.0334 | -- |  | 4.5 | 5
|  |  |  | 33.8837 | 6.8852 | 7 | 32.7 | 32.7 | 8.3618 | 7.8618 | -- |  | 4.6 | 5
|  |  |  | 33.9837 | 6.9055 | 7 | 32.7 | 32.7 | 8.1936 | 7.6936 | -- |  | 4.7 | 5
|  |  |  | 34.0837 | 6.9258 | 7 | 32.7 | 32.7 | 8.0287 | 7.5287 | -- |  | 4.8 | 5
|  |  |  | 34.1837 | 6.9461 | 7 | 32.7 | 32.7 | 7.8672 | 7.3672 | -- |  | 4.9 | 5
|  |  |  | 34.2837 | 6.9664 | 7 | 32.7 | 32.7 | 7.709 | 7.209 | -- |  | 5 | 5
|  |  |  | 34.3837 | 6.9868 | 7 | 32.7 | 32.7 | 7.5539 | 7.0539 | -- |  | 5.1 | 5
|  |  |  | 34.4837 | 7.0071 | 7 | 32.7 | 32.7 | 7.402 | 6.902 | -- |  | 5.2 | 5
|  |  |  | 34.5837 | 7.0274 | 7 | 32.7 | 32.7 | 7.2531 | 6.7531 | -- |  | 5.3 | 5
|  |  |  | 34.6837 | 7.0477 | 7 | 32.7 | 32.7 | 7.1072 | 6.6072 | -- |  | 5.4 | 5
|  |  |  | 34.7837 | 7.068 | 7 | 32.7 | 32.7 | 6.9642 | 6.4642 | -- |  | 5.5 | 6
|  |  |  | 34.8837 | 7.0884 | 7 | 32.7 | 32.7 | 6.8242 | 6.3242 | -- |  | 5.6 | 6
|  |  |  | 34.9837 | 7.1087 | 7 | 32.7 | 32.7 | 6.6869 | 6.1869 | -- |  | 5.7 | 6
|  |  |  | 35.0837 | 7.129 | 7 | 32.7 | 32.7 | 6.5524 | 6.0524 | -- |  | 5.8 | 6
|  |  |  | 35.1837 | 7.1493 | 7 | 32.7 | 32.7 | 6.4206 | 5.9206 | -- |  | 5.9 | 6
|  |  |  | 35.2837 | 7.1696 | 7 | 32.7 | 32.7 | 6.2914 | 5.7914 | -- |  | 6 | 6
|  |  |  | 35.3837 | 7.19 | 7 | 32.7 | 32.7 | 6.1649 | 5.6649 | -- |  | 6.1 | 6
|  |  |  | 35.4837 | 7.2103 | 7 | 32.7 | 32.7 | 6.0409 | 5.5409 | -- |  | 6.2 | 6
|  |  |  | 35.5837 | 7.2306 | 7 | 32.7 | 32.7 | 5.9194 | 5.4194 | -- |  | 6.3 | 6
|  |  |  | 35.6837 | 7.2509 | 7 | 32.7 | 32.7 | 5.8003 | 5.3003 | -- |  | 6.4 | 6
|  |  |  | 35.7837 | 7.2712 | 7 | 32.7 | 32.7 | 5.6836 | 5.1836 | -- |  | 6.5 | 7
|  |  |  | 35.8837 | 7.2916 | 7 | 32.7 | 32.7 | 5.5693 | 5.0693 | -- |  | 6.6 | 7
|  |  |  | 35.9837 | 7.3119 | 7 | 32.7 | 32.7 | 5.4573 | 4.9573 | -- |  | 6.7 | 7
|  |  |  | 36.0837 | 7.3322 | 7 | 32.7 | 32.7 | 5.3475 | 4.8475 | -- |  | 6.8 | 7
|  |  |  | 36.1837 | 7.3525 | 7 | 32.7 | 32.7 | 5.2399 | 4.7399 | -- |  | 6.9 | 7
|  |  |  | 36.2837 | 7.3728 | 7 | 32.7 | 32.7 | 5.1345 | 4.6345 | -- |  | 7 | 7
|  |  |  | 36.3837 | 7.3932 | 7 | 32.7 | 32.7 | 5.0312 | 4.5312 | -- |  | 7.1 | 7
|  |  |  | 36.4837 | 7.4135 | 7 | 32.7 | 32.7 | 4.93 | 4.43 | -- |  | 7.2 | 7
|  |  |  | 36.5837 | 7.4338 | 7 | 32.7 | 32.7 | 4.8309 | 4.3309 | -- |  | 7.3 | 7
|  |  |  | 36.6837 | 7.4541 | 7 | 32.7 | 32.7 | 4.7337 | 4.2337 | -- |  | 7.4 | 7
|  |  |  | 36.7837 | 7.4744 | 7 | 32.7 | 32.7 | 4.6385 | 4.1385 | -- |  | 7.5 | 8
|  |  |  | 36.8837 | 7.4948 | 7 | 32.7 | 32.7 | 4.5452 | 4.0452 | -- |  | 7.6 | 8
|  |  |  | 36.9837 | 7.5151 | 8 | 40.6 | 40.6 | 5.5297 | 5.0297 | -- |  | 7.7 | 8
|  |  |  | 37.0837 | 7.5354 | 8 | 40.6 | 40.6 | 5.4185 | 4.9185 | -- |  | 7.8 | 8
|  |  |  | 37.1837 | 7.5557 | 8 | 40.6 | 40.6 | 5.3095 | 4.8095 | -- |  | 7.9 | 8
|  |  |  | 37.2837 | 7.576 | 8 | 40.6 | 40.6 | 5.2027 | 4.7027 | -- |  | 8 | 8
|  |  |  | 37.3837 | 7.5964 | 8 | 40.6 | 40.6 | 5.0981 | 4.5981 | -- |  | 8.1 | 8
|  |  |  | 37.4837 | 7.6167 | 8 | 40.6 | 40.6 | 4.9955 | 4.4955 | -- |  | 8.2 | 8
|  |  |  | 37.5837 | 7.637 | 8 | 40.6 | 40.6 | 4.895 | 4.395 | -- |  | 8.3 | 8
|  |  |  | 37.6837 | 7.6573 | 8 | 40.6 | 40.6 | 4.7966 | 4.2966 | -- |  | 8.4 | 8
|  |  |  | 37.7837 | 7.6776 | 8 | 40.6 | 40.6 | 4.7001 | 4.2001 | -- |  | 8.5 | 8
|  |  |  | 37.8837 | 7.698 | 8 | 40.6 | 40.6 | 4.6055 | 4.1055 | -- |  | 8.6 | 8
|  |  |  | 37.9837 | 7.7183 | 8 | 40.6 | 40.6 | 4.5129 | 4.0129 | -- |  | 8.7 | 8
|  |  |  | 38.0837 | 7.7386 | 8 | 40.6 | 40.6 | 4.4221 | 3.9221 | -- |  | 8.8 | 8
|  |  |  | 38.1837 | 7.7589 | 8 | 40.6 | 40.6 | 4.3332 | 3.8332 | -- |  | 8.9 | 8
|  |  |  | 38.2837 | 7.7792 | 8 | 40.6 | 40.6 | 4.246 | 3.746 | -- |  | 9 | 10
|  |  |  | 38.3837 | 7.7996 | 8 | 40.6 | 40.6 | 4.1606 | 3.6606 | -- |  | 9.1 | 10
|  |  |  | 38.4837 | 7.8199 | 8 | 40.6 | 40.6 | 4.0769 | 3.5769 | -- |  | 9.2 | 10
|  |  |  | 38.5837 | 7.8402 | 8 | 40.6 | 40.6 | 3.9949 | 3.4949 | -- |  | 9.3 | 10
|  |  |  | 38.6837 | 7.8605 | 8 | 40.6 | 40.6 | 3.9146 | 3.4146 | -- |  | 9.4 | 10
|  |  |  | 38.7837 | 7.8808 | 8 | 40.6 | 40.6 | 3.8358 | 3.3358 | -- |  | 9.5 | 10
|  |  |  | 38.8837 | 7.9012 | 8 | 40.6 | 40.6 | 3.7587 | 3.2587 | -- |  | 9.6 | 10
|  |  |  | 38.9837 | 7.9215 | 8 | 40.6 | 40.6 | 3.683 | 3.183 | -- |  | 9.7 | 10
|  |  |  | 39.0837 | 7.9418 | 8 | 40.6 | 40.6 | 3.609 | 3.109 | -- |  | 9.8 | 10
|  |  |  | 39.1837 | 7.9621 | 8 | 40.6 | 40.6 | 3.5364 | 3.0364 | -- |  | 9.9 | 10
|  |  |  | 39.2837 | 7.9824 | 8 | 40.6 | 40.6 | 3.4652 | 2.9652 | -- |  | 10 | 10
|  |  |  | 39.3837 | 8.0028 | 8 | 40.6 | 40.6 | 3.3955 | 2.8955 | -- |  | 10.1 | 10
|  |  |  | 39.4837 | 8.0231 | 8 | 40.6 | 40.6 | 3.3272 | 2.8272 | -- |  | 10.2 | 10
|  |  |  | 39.5837 | 8.0434 | 8 | 40.6 | 40.6 | 3.2603 | 2.7603 | -- |  | 10.3 | 10
|  |  |  | 39.6837 | 8.0637 | 8 | 40.6 | 40.6 | 3.1947 | 2.6947 | -- |  | 10.4 | 10
|  |  |  | 39.7837 | 8.084 | 8 | 40.6 | 40.6 | 3.1305 | 2.6305 | -- |  | 10.5 | 10
|  |  |  | 39.8837 | 8.1044 | 8 | 40.6 | 40.6 | 3.0675 | 2.5675 | -- |  | 10.6 | 10
|  |  |  | 39.9837 | 8.1247 | 8 | 40.6 | 40.6 | 3.0058 | 2.5058 | -- |  | 10.7 | 10
|  |  |  | 40.0837 | 8.145 | 8 | 40.6 | 40.6 | 2.9453 | 2.4453 | -- |  | 10.8 | 10
|  |  |  | 40.1837 | 8.1653 | 8 | 40.6 | 40.6 | 2.8861 | 2.3861 | -- |  | 10.9 | 10
|  |  |  | 40.2837 | 8.1856 | 8 | 40.6 | 40.6 | 2.828 | 2.328 | -- |  | 11 | 10
|  |  |  | 40.3837 | 8.206 | 8 | 40.6 | 40.6 | 2.7711 | 2.2711 | -- |  | 11.1 | 10
|  |  |  | 40.4837 | 8.2263 | 8 | 40.6 | 40.6 | 2.7154 | 2.2154 | -- |  | 11.2 | 10
|  |  |  | 40.5837 | 8.2466 | 8 | 40.6 | 40.6 | 2.6608 | 2.1608 | -- |  | 11.3 | 10
|  |  |  | 40.6837 | 8.2669 | 8 | 40.6 | 40.6 | 2.6073 | 2.1073 | -- |  | 11.4 | 10
|  |  |  | 40.7837 | 8.2872 | 8 | 40.6 | 40.6 | 2.5548 | 2.0548 | -- |  | 11.5 | 10
|  |  |  | 40.8837 | 8.3076 | 8 | 40.6 | 40.6 | 2.5034 | 2.0034 | -- |  | 11.6 | 10
|  |  |  | 40.9837 | 8.3279 | 8 | 40.6 | 40.6 | 2.4531 | 1.9531 | -- |  | 11.7 | 10
|  |  |  | 41.0837 | 8.3482 | 8 | 40.6 | 40.6 | 2.4037 | 1.9037 | -- |  | 11.8 | 10
|  |  |  | 41.1837 | 8.3685 | 8 | 40.6 | 40.6 | 2.3554 | 1.8554 | -- |  | 11.9 | 10
|  |  |  | 41.2837 | 8.3888 | 8 | 40.6 | 40.6 | 2.308 | 1.808 | -- |  | 12 | 10
|  |  |  | 41.3837 | 8.4092 | 8 | 40.6 | 40.6 | 2.2616 | 1.7616 | -- |  | 12.1 | 10
|  |  |  | 41.4837 | 8.4295 | 8 | 40.6 | 40.6 | 2.2161 | 1.7161 | -- |  | 12.2 | 10
|  |  |  | 41.5837 | 8.4498 | 8 | 40.6 | 40.6 | 2.1715 | 1.6715 | -- |  | 12.3 | 10
|  |  |  | 41.6837 | 8.4701 | 8 | 40.6 | 40.6 | 2.1278 | 1.6278 | -- |  | 12.4 | 10
|  |  |  | 41.7837 | 8.4904 | 8 | 40.6 | 40.6 | 2.085 | 1.585 | -- |  | 12.5 | 15
|  |  |  | 41.8837 | 8.5108 | 8 | 40.6 | 40.6 | 2.0431 | 1.5431 | -- |  | 12.6 | 15
|  |  |  | 41.9837 | 8.5311 | 8 | 40.6 | 40.6 | 2.002 | 1.502 | -- |  | 12.7 | 15
|  |  |  | 42.0837 | 8.5514 | 8 | 40.6 | 40.6 | 1.9617 | 1.4617 | -- |  | 12.8 | 15
|  |  |  | 42.1837 | 8.5717 | 8 | 40.6 | 40.6 | 1.9223 | 1.4223 | -- |  | 12.9 | 15
|  |  |  | 42.2837 | 8.592 | 8 | 40.6 | 40.6 | 1.8836 | 1.3836 | -- |  | 13 | 15
|  |  |  | 42.3837 | 8.6124 | 8 | 40.6 | 40.6 | 1.8457 | 1.3457 | -- |  | 13.1 | 15
|  |  |  | 42.4837 | 8.6327 | 8 | 40.6 | 40.6 | 1.8086 | 1.3086 | -- |  | 13.2 | 15
|  |  |  | 42.5837 | 8.653 | 8 | 40.6 | 40.6 | 1.7722 | 1.2722 | -- |  | 13.3 | 15
|  |  |  | 42.6837 | 8.6733 | 8 | 40.6 | 40.6 | 1.7366 | 1.2366 | -- |  | 13.4 | 15
|  |  |  | 42.7837 | 8.6936 | 8 | 40.6 | 40.6 | 1.7016 | 1.2016 | -- |  | 13.5 | 15
|  |  |  | 42.8837 | 8.714 | 8 | 40.6 | 40.6 | 1.6674 | 1.1674 | -- |  | 13.6 | 15
|  |  |  | 42.9837 | 8.7343 | 8 | 40.6 | 40.6 | 1.6339 | 1.1339 | -- |  | 13.7 | 15
|  |  |  | 43.0837 | 8.7546 | 8 | 40.6 | 40.6 | 1.601 | 1.101 | -- |  | 13.8 | 15
|  |  |  | 43.1837 | 8.7749 | 8 | 40.6 | 40.6 | 1.5688 | 1.0688 | -- |  | 13.9 | 15
|  |  |  | 43.2837 | 8.7952 | 8 | 40.6 | 40.6 | 1.5372 | 1.0372 | -- |  | 14 | 15
|  |  |  | 43.3837 | 8.8156 | 8 | 40.6 | 40.6 | 1.5063 | 1.0063 | -- |  | 14.1 | 15
|  |  |  | 43.4837 | 8.8359 | 8 | 40.6 | 40.6 | 1.476 | 0.97601 | -- |  | 14.2 | 15
|  |  |  | 43.5837 | 8.8562 | 8 | 40.6 | 40.6 | 1.4463 | 0.94632 | -- |  | 14.3 | 15
|  |  |  | 43.6837 | 8.8765 | 8 | 40.6 | 40.6 | 1.4172 | 0.91723 | -- |  | 14.4 | 15
|  |  |  | 43.7837 | 8.8968 | 8 | 40.6 | 40.6 | 1.3887 | 0.88872 | -- |  | 14.5 | 15
|  |  |  | 43.8837 | 8.9172 | 8 | 40.6 | 40.6 | 1.3608 | 0.86078 | -- |  | 14.6 | 15
|  |  |  | 43.9837 | 8.9375 | 8 | 40.6 | 40.6 | 1.3334 | 0.83341 | -- |  | 14.7 | 15
|  |  |  | 44.0837 | 8.9578 | 8 | 40.6 | 40.6 | 1.3066 | 0.80659 | -- |  | 14.8 | 15
|  |  |  | 44.1837 | 8.9781 | 8 | 40.6 | 40.6 | 1.2803 | 0.78031 | -- |  | 14.9 | 15
|  |  |  | 44.2837 | 8.9984 | 8 | 40.6 | 40.6 | 1.2546 | 0.75455 | -- |  | 15 | 15
|  |  |  | 44.3837 | 9.0188 | 10 | 59 | 59 | 1.7864 | 1.2864 | -- |  | 15.1 | 15
|  |  |  | 44.4837 | 9.0391 | 10 | 59 | 59 | 1.7505 | 1.2505 | -- |  | 15.2 | 15
|  |  |  | 44.5837 | 9.0594 | 10 | 59 | 59 | 1.7153 | 1.2153 | -- |  | 15.3 | 15
|  |  |  | 44.6837 | 9.0797 | 10 | 59 | 59 | 1.6808 | 1.1808 | -- |  | 15.4 | 15
|  |  |  | 44.7837 | 9.1 | 10 | 59 | 59 | 1.647 | 1.147 | -- |  | 15.5 | 15
|  |  |  | 44.8837 | 9.1204 | 10 | 59 | 59 | 1.6139 | 1.1139 | -- |  | 15.6 | 15
|  |  |  | 44.9837 | 9.1407 | 10 | 59 | 59 | 1.5814 | 1.0814 | -- |  | 15.7 | 15
|  |  |  | 45.0837 | 9.161 | 10 | 59 | 59 | 1.5496 | 1.0496 | -- |  | 15.8 | 15
|  |  |  | 45.1837 | 9.1813 | 10 | 59 | 59 | 1.5184 | 1.0184 | -- |  | 15.9 | 15
|  |  |  | 45.2837 | 9.2016 | 10 | 59 | 59 | 1.4879 | 0.98788 | -- |  | 16 | 15
|  |  |  | 45.3837 | 9.222 | 10 | 59 | 59 | 1.4579 | 0.95795 | -- |  | 16.1 | 15
|  |  |  | 45.4837 | 9.2423 | 10 | 59 | 59 | 1.4286 | 0.92862 | -- |  | 16.2 | 15
|  |  |  | 45.5837 | 9.2626 | 10 | 59 | 59 | 1.3999 | 0.89989 | -- |  | 16.3 | 15
|  |  |  | 45.6837 | 9.2829 | 10 | 59 | 59 | 1.3717 | 0.87173 | -- |  | 16.4 | 15
|  |  |  | 45.7837 | 9.3032 | 10 | 59 | 59 | 1.3441 | 0.84413 | -- |  | 16.5 | 15
|  |  |  | 45.8837 | 9.3236 | 10 | 59 | 59 | 1.3171 | 0.8171 | -- |  | 16.6 | 15
|  |  |  | 45.9837 | 9.3439 | 10 | 59 | 59 | 1.2906 | 0.7906 | -- |  | 16.7 | 15
|  |  |  | 46.0837 | 9.3642 | 10 | 59 | 59 | 1.2646 | 0.76464 | -- |  | 16.8 | 15
|  |  |  | 46.1837 | 9.3845 | 10 | 59 | 59 | 1.2392 | 0.73921 | -- |  | 16.9 | 15
|  |  |  | 46.2837 | 9.4048 | 10 | 59 | 59 | 1.2143 | 0.71428 | -- |  | 17 | 15
|  |  |  | 46.3837 | 9.4252 | 10 | 59 | 59 | 1.1899 | 0.68985 | -- |  | 17.1 | 15
|  |  |  | 46.4837 | 9.4455 | 10 | 59 | 59 | 1.1659 | 0.66592 | -- |  | 17.2 | 15
|  |  |  | 46.5837 | 9.4658 | 10 | 59 | 59 | 1.1425 | 0.64247 | -- |  | 17.3 | 15
|  |  |  | 46.6837 | 9.4861 | 10 | 59 | 59 | 1.1195 | 0.61949 | -- |  | 17.4 | 15
|  |  |  | 46.7837 | 9.5064 | 10 | 59 | 59 | 1.097 | 0.59697 | -- |  | 17.5 | 20
|  |  |  | 46.8837 | 9.5268 | 10 | 59 | 59 | 1.0749 | 0.5749 | -- |  | 17.6 | 20
|  |  |  | 46.9837 | 9.5471 | 10 | 59 | 59 | 1.0533 | 0.55328 | -- |  | 17.7 | 20
|  |  |  | 47.0837 | 9.5674 | 10 | 59 | 59 | 1.0321 | 0.53209 | -- |  | 17.8 | 20
|  |  |  | 47.1837 | 9.5877 | 10 | 59 | 59 | 1.0113 | 0.51133 | -- |  | 17.9 | 20
|  |  |  | 47.2837 | 9.608 | 10 | 59 | 59 | 0.99099 | 0.49099 | -- |  | 18 | 20
|  |  |  | 47.3837 | 9.6284 | 10 | 59 | 59 | 0.97106 | 0.47106 | -- |  | 18.1 | 20
|  |  |  | 47.4837 | 9.6487 | 10 | 59 | 59 | 0.95152 | 0.45152 | -- |  | 18.2 | 20
|  |  |  | 47.5837 | 9.669 | 10 | 59 | 59 | 0.93238 | 0.43238 | -- |  | 18.3 | 20
|  |  |  | 47.6837 | 9.6893 | 10 | 59 | 59 | 0.91363 | 0.41363 | -- |  | 18.4 | 20
|  |  |  | 47.7837 | 9.7096 | 10 | 59 | 59 | 0.89525 | 0.39525 | -- |  | 18.5 | 20
|  |  |  | 47.8837 | 9.73 | 10 | 59 | 59 | 0.87724 | 0.37724 | -- |  | 18.6 | 20
|  |  |  | 47.9837 | 9.7503 | 10 | 59 | 59 | 0.8596 | 0.3596 | -- |  | 18.7 | 20
|  |  |  | 48.0837 | 9.7706 | 10 | 59 | 59 | 0.84231 | 0.34231 | -- |  | 18.8 | 20
|  |  |  | 48.1837 | 9.7909 | 10 | 59 | 59 | 0.82536 | 0.32536 | -- |  | 18.9 | 20
|  |  |  | 48.2837 | 9.8112 | 10 | 59 | 59 | 0.80876 | 0.30876 | -- |  | 19 | 20
|  |  |  | 48.3837 | 9.8316 | 10 | 59 | 59 | 0.79249 | 0.29249 | -- |  | 19.1 | 20
|  |  |  | 48.4837 | 9.8519 | 10 | 59 | 59 | 0.77655 | 0.27655 | -- |  | 19.2 | 20
|  |  |  | 48.5837 | 9.8722 | 10 | 59 | 59 | 0.76093 | 0.26093 | -- |  | 19.3 | 20
|  |  |  | 48.6837 | 9.8925 | 10 | 59 | 59 | 0.74563 | 0.24563 | -- |  | 19.4 | 20
|  |  |  | 48.7837 | 9.9128 | 10 | 59 | 59 | 0.73063 | 0.23063 | -- |  | 19.5 | 20
|  |  |  | 48.8837 | 9.9332 | 10 | 59 | 59 | 0.71593 | 0.21593 | -- |  | 19.6 | 20
|  |  |  | 48.9837 | 9.9535 | 10 | 59 | 59 | 0.70153 | 0.20153 | -- |  | 19.7 | 20
|  |  |  | 49.0837 | 9.9738 | 10 | 59 | 59 | 0.68742 | 0.18742 | -- |  | 19.8 | 20
|  |  |  | 49.1837 | 9.9941 | 10 | 59 | 59 | 0.67359 | 0.17359 | -- |  | 19.9 | 20
|  |  |  | 49.2837 | 10.0144 | 10 | 59 | 59 | 0.66004 | 0.16004 | -- |  | 20 | 20
|  |  |  | 49.3837 | 10.0348 | 10 | 59 | 59 | 0.64677 | 0.14677 | -- |  | 20.1 | 20
|  |  |  | 49.4837 | 10.0551 | 10 | 59 | 59 | 0.63376 | 0.13376 | -- |  | 20.2 | 20
|  |  |  | 49.5837 | 10.0754 | 10 | 59 | 59 | 0.62101 | 0.12101 | -- |  | 20.3 | 20
|  |  |  | 49.6837 | 10.0957 | 10 | 59 | 59 | 0.60852 | 0.10852 | -- |  | 20.4 | 20
|  |  |  | 49.7837 | 10.116 | 10 | 59 | 59 | 0.59628 | 0.096277 | -- |  | 20.5 | 20
|  |  |  | 49.8837 | 10.1364 | 10 | 59 | 59 | 0.58428 | 0.084283 | -- |  | 20.6 | 20
|  |  |  | 49.9837 | 10.1567 | 10 | 59 | 59 | 0.57253 | 0.07253 | -- |  | 20.7 | 20
|  |  |  | 50.0837 | 10.177 | 10 | 59 | 59 | 0.56101 | 0.061014 | -- |  | 20.8 | 20
|  |  |  | 50.1837 | 10.1973 | 10 | 59 | 59 | 0.54973 | 0.049729 | -- |  | 20.9 | 20
|  |  |  | 50.2837 | 10.2176 | 10 | 59 | 59 | 0.53867 | 0.038671 | -- |  | 21 | 20
|  |  |  | 50.3837 | 10.238 | 10 | 59 | 59 | 0.52784 | 0.027836 | -- |  | 21.1 | 20
|  |  |  | 50.4837 | 10.2583 | 10 | 59 | 59 | 0.51722 | 0.017219 | -- |  | 21.2 | 20
|  |  |  | 50.5837 | 10.2786 | 10 | 59 | 59 | 0.50681 | 0.0068147 | -- |  | 21.3 | 20
|  |  |  | 50.6837 | 10.2989 | 10 | 59 | 59 | 0.49662 | -0.0033798 | 50.6837 |  | 21.4 | 20
|  |  |  | 50.7837 | 10.3192 | 10 | 59 | 59 | 0.48663 | -0.013369 | 50.7837 |  | 21.5 | 20
|  |  |  | 50.8837 | 10.3396 | 10 | 59 | 59 | 0.47684 | -0.023158 | 50.8837 |  | 21.6 | 20
|  |  |  | 50.9837 | 10.3599 | 10 | 59 | 59 | 0.46725 | -0.032749 | 50.9837 |  | 21.7 | 20
|  |  |  | 51.0837 | 10.3802 | 10 | 59 | 59 | 0.45785 | -0.042148 | 51.0837 |  | 21.8 | 20
|  |  |  | 51.1837 | 10.4005 | 10 | 59 | 59 | 0.44864 | -0.051358 | 51.1837 |  | 21.9 | 20
|  |  |  | 51.2837 | 10.4208 | 10 | 59 | 59 | 0.43962 | -0.060382 | 51.2837 |  | 22 | 20
|  |  |  | 51.3837 | 10.4412 | 10 | 59 | 59 | 0.43077 | -0.069225 | 51.3837 |  | 22.1 | 20
|  |  |  | 51.4837 | 10.4615 | 10 | 59 | 59 | 0.42211 | -0.07789 | 51.4837 |  | 22.2 | 20
|  |  |  | 51.5837 | 10.4818 | 10 | 59 | 59 | 0.41362 | -0.086381 | 51.5837 |  | 22.3 | 20
|  |  |  | 51.6837 | 10.5021 | 10 | 59 | 59 | 0.4053 | -0.094701 | 51.6837 |  | 22.4 | 20
|  |  |  | 51.7837 | 10.5224 | 10 | 59 | 59 | 0.39715 | -0.10285 | 51.7837 |  | 22.5 | 25
|  |  |  | 51.8837 | 10.5428 | 10 | 59 | 59 | 0.38916 | -0.11084 | 51.8837 |  | 22.6 | 25
|  |  |  | 51.9837 | 10.5631 | 10 | 59 | 59 | 0.38133 | -0.11867 | 51.9837 |  | 22.7 | 25
|  |  |  | 52.0837 | 10.5834 | 10 | 59 | 59 | 0.37366 | -0.12634 | 52.0837 |  | 22.8 | 25
|  |  |  | 52.1837 | 10.6037 | 10 | 59 | 59 | 0.36614 | -0.13386 | 52.1837 |  | 22.9 | 25
|  |  |  | 52.2837 | 10.624 | 10 | 59 | 59 | 0.35878 | -0.14122 | 52.2837 |  | 23 | 25
|  |  |  | 52.3837 | 10.6444 | 10 | 59 | 59 | 0.35156 | -0.14844 | 52.3837 |  | 23.1 | 25
|  |  |  | 52.4837 | 10.6647 | 10 | 59 | 59 | 0.34449 | -0.15551 | 52.4837 |  | 23.2 | 25
|  |  |  | 52.5837 | 10.685 | 10 | 59 | 59 | 0.33756 | -0.16244 | 52.5837 |  | 23.3 | 25
|  |  |  | 52.6837 | 10.7053 | 10 | 59 | 59 | 0.33077 | -0.16923 | 52.6837 |  | 23.4 | 25
|  |  |  | 52.7837 | 10.7256 | 10 | 59 | 59 | 0.32412 | -0.17588 | 52.7837 |  | 23.5 | 25
|  |  |  | 52.8837 | 10.746 | 10 | 59 | 59 | 0.3176 | -0.1824 | 52.8837 |  | 23.6 | 25
|  |  |  | 52.9837 | 10.7663 | 10 | 59 | 59 | 0.31121 | -0.18879 | 52.9837 |  | 23.7 | 25
|  |  |  | 53.0837 | 10.7866 | 10 | 59 | 59 | 0.30495 | -0.19505 | 53.0837 |  | 23.8 | 25
|  |  |  | 53.1837 | 10.8069 | 10 | 59 | 59 | 0.29882 | -0.20118 | 53.1837 |  | 23.9 | 25
|  |  |  | 53.2837 | 10.8272 | 10 | 59 | 59 | 0.2928 | -0.2072 | 53.2837 |  | 24 | 25
|  |  |  | 53.3837 | 10.8476 | 10 | 59 | 59 | 0.28691 | -0.21309 | 53.3837 |  | 24.1 | 25
|  |  |  | 53.4837 | 10.8679 | 10 | 59 | 59 | 0.28114 | -0.21886 | 53.4837 |  | 24.2 | 25
|  |  |  | 53.5837 | 10.8882 | 10 | 59 | 59 | 0.27549 | -0.22451 | 53.5837 |  | 24.3 | 25
|  |  |  | 53.6837 | 10.9085 | 10 | 59 | 59 | 0.26995 | -0.23005 | 53.6837 |  | 24.4 | 25
|  |  |  | 53.7837 | 10.9288 | 10 | 59 | 59 | 0.26452 | -0.23548 | 53.7837 |  | 24.5 | 25
|  |  |  | 53.8837 | 10.9492 | 10 | 59 | 59 | 0.2592 | -0.2408 | 53.8837 |  | 24.6 | 25
|  |  |  | 53.9837 | 10.9695 | 10 | 59 | 59 | 0.25398 | -0.24602 | 53.9837 |  | 24.7 | 25
|  |  |  | 54.0837 | 10.9898 | 10 | 59 | 59 | 0.24887 | -0.25113 | 54.0837 |  | 24.8 | 25
|  |  |  | 54.1837 | 11.0101 | 10 | 59 | 59 | 0.24387 | -0.25613 | 54.1837 |  | 24.9 | 25
|  |  |  | 54.2837 | 11.0304 | 10 | 59 | 59 | 0.23896 | -0.26104 | 54.2837 |  | 25 | 25
|  |  |  | 54.3837 | 11.0508 | 10 | 59 | 59 | 0.23416 | -0.26584 | 54.3837 |  | 25.1 | 25
|  |  |  | 54.4837 | 11.0711 | 10 | 59 | 59 | 0.22945 | -0.27055 | 54.4837 |  | 25.2 | 25
|  |  |  | 54.5837 | 11.0914 | 10 | 59 | 59 | 0.22483 | -0.27517 | 54.5837 |  | 25.3 | 25
|  |  |  | 54.6837 | 11.1117 | 10 | 59 | 59 | 0.22031 | -0.27969 | 54.6837 |  | 25.4 | 25
|  |  |  | 54.7837 | 11.132 | 10 | 59 | 59 | 0.21588 | -0.28412 | 54.7837 |  | 25.5 | 25
|  |  |  | 54.8837 | 11.1524 | 10 | 59 | 59 | 0.21153 | -0.28847 | 54.8837 |  | 25.6 | 25
|  |  |  | 54.9837 | 11.1727 | 10 | 59 | 59 | 0.20728 | -0.29272 | 54.9837 |  | 25.7 | 25
|  |  |  | 55.0837 | 11.193 | 10 | 59 | 59 | 0.20311 | -0.29689 | 55.0837 |  | 25.8 | 25
|  |  |  | 55.1837 | 11.2133 | 10 | 59 | 59 | 0.19902 | -0.30098 | 55.1837 |  | 25.9 | 25
|  |  |  | 55.2837 | 11.2336 | 10 | 59 | 59 | 0.19502 | -0.30498 | 55.2837 |  | 26 | 25
|  |  |  | 55.3837 | 11.254 | 10 | 59 | 59 | 0.1911 | -0.3089 | 55.3837 |  | 26.1 | 25
|  |  |  | 55.4837 | 11.2743 | 10 | 59 | 59 | 0.18725 | -0.31275 | 55.4837 |  | 26.2 | 25
|  |  |  | 55.5837 | 11.2946 | 10 | 59 | 59 | 0.18349 | -0.31651 | 55.5837 |  | 26.3 | 25
|  |  |  | 55.6837 | 11.3149 | 10 | 59 | 59 | 0.1798 | -0.3202 | 55.6837 |  | 26.4 | 25
|  |  |  | 55.7837 | 11.3352 | 10 | 59 | 59 | 0.17618 | -0.32382 | 55.7837 |  | 26.5 | 25
|  |  |  | 55.8837 | 11.3556 | 10 | 59 | 59 | 0.17264 | -0.32736 | 55.8837 |  | 26.6 | 25
|  |  |  | 55.9837 | 11.3759 | 10 | 59 | 59 | 0.16916 | -0.33084 | 55.9837 |  | 26.7 | 25
|  |  |  | 56.0837 | 11.3962 | 10 | 59 | 59 | 0.16576 | -0.33424 | 56.0837 |  | 26.8 | 25
|  |  |  | 56.1837 | 11.4165 | 10 | 59 | 59 | 0.16243 | -0.33757 | 56.1837 |  | 26.9 | 25
|  |  |  | 56.2837 | 11.4368 | 10 | 59 | 59 | 0.15916 | -0.34084 | 56.2837 |  | 27 | 25
|  |  |  | 56.3837 | 11.4572 | 10 | 59 | 59 | 0.15596 | -0.34404 | 56.3837 |  | 27.1 | 25
|  |  |  | 56.4837 | 11.4775 | 10 | 59 | 59 | 0.15282 | -0.34718 | 56.4837 |  | 27.2 | 25
|  |  |  | 56.5837 | 11.4978 | 10 | 59 | 59 | 0.14975 | -0.35025 | 56.5837 |  | 27.3 | 25
|  |  |  | 56.6837 | 11.5181 | 10 | 59 | 59 | 0.14673 | -0.35327 | 56.6837 |  | 27.4 | 25
|  |  |  | 56.7837 | 11.5384 | 10 | 59 | 59 | 0.14378 | -0.35622 | 56.7837 |  | 27.5 | 30
|  |  |  | 56.8837 | 11.5588 | 10 | 59 | 59 | 0.14089 | -0.35911 | 56.8837 |  | 27.6 | 30
|  |  |  | 56.9837 | 11.5791 | 10 | 59 | 59 | 0.13806 | -0.36194 | 56.9837 |  | 27.7 | 30
|  |  |  | 57.0837 | 11.5994 | 10 | 59 | 59 | 0.13528 | -0.36472 | 57.0837 |  | 27.8 | 30
|  |  |  | 57.1837 | 11.6197 | 10 | 59 | 59 | 0.13256 | -0.36744 | 57.1837 |  | 27.9 | 30
|  |  |  | 57.2837 | 11.64 | 10 | 59 | 59 | 0.12989 | -0.37011 | 57.2837 |  | 28 | 30
|  |  |  | 57.3837 | 11.6604 | 10 | 59 | 59 | 0.12728 | -0.37272 | 57.3837 |  | 28.1 | 30
|  |  |  | 57.4837 | 11.6807 | 10 | 59 | 59 | 0.12472 | -0.37528 | 57.4837 |  | 28.2 | 30
|  |  |  | 57.5837 | 11.701 | 10 | 59 | 59 | 0.12221 | -0.37779 | 57.5837 |  | 28.3 | 30
|  |  |  | 57.6837 | 11.7213 | 10 | 59 | 59 | 0.11975 | -0.38025 | 57.6837 |  | 28.4 | 30
|  |  |  | 57.7837 | 11.7416 | 10 | 59 | 59 | 0.11734 | -0.38266 | 57.7837 |  | 28.5 | 30
|  |  |  | 57.8837 | 11.762 | 10 | 59 | 59 | 0.11498 | -0.38502 | 57.8837 |  | 28.6 | 30
|  |  |  | 57.9837 | 11.7823 | 10 | 59 | 59 | 0.11267 | -0.38733 | 57.9837 |  | 28.7 | 30
|  |  |  | 58.0837 | 11.8026 | 10 | 59 | 59 | 0.1104 | -0.3896 | 58.0837 |  | 28.8 | 30
|  |  |  | 58.1837 | 11.8229 | 10 | 59 | 59 | 0.10818 | -0.39182 | 58.1837 |  | 28.9 | 30
|  |  |  | 58.2837 | 11.8432 | 10 | 59 | 59 | 0.10601 | -0.39399 | 58.2837 |  | 29 | 30
|  |  |  | 58.3837 | 11.8636 | 10 | 59 | 59 | 0.10387 | -0.39613 | 58.3837 |  | 29.1 | 30
|  |  |  | 58.4837 | 11.8839 | 10 | 59 | 59 | 0.10179 | -0.39821 | 58.4837 |  | 29.2 | 30
|  |  |  | 58.5837 | 11.9042 | 10 | 59 | 59 | 0.099738 | -0.40026 | 58.5837 |  | 29.3 | 30
|  |  |  | 58.6837 | 11.9245 | 10 | 59 | 59 | 0.097732 | -0.40227 | 58.6837 |  | 29.4 | 30
|  |  |  | 58.7837 | 11.9448 | 10 | 59 | 59 | 0.095766 | -0.40423 | 58.7837 |  | 29.5 | 30
|  |  |  | 58.8837 | 11.9652 | 10 | 59 | 59 | 0.093839 | -0.40616 | 58.8837 |  | 29.6 | 30
|  |  |  | 58.9837 | 11.9855 | 10 | 59 | 59 | 0.091952 | -0.40805 | 58.9837 |  | 29.7 | 30
|  |  |  | 59.0837 | 12.0058 | 10 | 59 | 59 | 0.090102 | -0.4099 | 59.0837 |  | 29.8 | 30
|  |  |  | 59.1837 | 12.0261 | 10 | 59 | 59 | 0.08829 | -0.41171 | 59.1837 |  | 29.9 | 30
|  |  |  | 59.2837 | 12.0464 | 10 | 59 | 59 | 0.086514 | -0.41349 | 59.2837 |  | 30 | 30
|  |  |  | 59.3837 | 12.0668 | 10 | 59 | 59 | 0.084774 | -0.41523 | 59.3837 |  | 30.1 | 30
|  |  |  | 59.4837 | 12.0871 | 10 | 59 | 59 | 0.083068 | -0.41693 | 59.4837 |  | 30.2 | 30
|  |  |  | 59.5837 | 12.1074 | 10 | 59 | 59 | 0.081398 | -0.4186 | 59.5837 |  | 30.3 | 30
|  |  |  | 59.6837 | 12.1277 | 10 | 59 | 59 | 0.07976 | -0.42024 | 59.6837 |  | 30.4 | 30
|  |  |  | 59.7837 | 12.148 | 10 | 59 | 59 | 0.078156 | -0.42184 | 59.7837 |  | 30.5 | 30
|  |  |  | 59.8837 | 12.1684 | 10 | 59 | 59 | 0.076584 | -0.42342 | 59.8837 |  | 30.6 | 30
|  |  |  | 59.9837 | 12.1887 | 10 | 59 | 59 | 0.075043 | -0.42496 | 59.9837 |  | 30.7 | 30
|  |  |  | 60.0837 | 12.209 | 10 | 59 | 59 | 0.073534 | -0.42647 | 60.0837 |  | 30.8 | 30
|  |  |  | 60.1837 | 12.2293 | 10 | 59 | 59 | 0.072055 | -0.42795 | 60.1837 |  | 30.9 | 30
|  |  |  | 60.2837 | 12.2496 | 10 | 59 | 59 | 0.070605 | -0.42939 | 60.2837 |  | 31 | 30
|  |  |  | 60.3837 | 12.27 | 10 | 59 | 59 | 0.069185 | -0.43081 | 60.3837 |  | 31.1 | 30
|  |  |  | 60.4837 | 12.2903 | 10 | 59 | 59 | 0.067793 | -0.43221 | 60.4837 |  | 31.2 | 30
|  |  |  | 60.5837 | 12.3106 | 10 | 59 | 59 | 0.06643 | -0.43357 | 60.5837 |  | 31.3 | 30
|  |  |  | 60.6837 | 12.3309 | 10 | 59 | 59 | 0.065094 | -0.43491 | 60.6837 |  | 31.4 | 30
|  |  |  | 60.7837 | 12.3512 | 10 | 59 | 59 | 0.063784 | -0.43622 | 60.7837 |  | 31.5 | 30
|  |  |  | 60.8837 | 12.3716 | 10 | 59 | 59 | 0.062501 | -0.4375 | 60.8837 |  | 31.6 | 30
|  |  |  | 60.9837 | 12.3919 | 10 | 59 | 59 | 0.061244 | -0.43876 | 60.9837 |  | 31.7 | 30
|  |  |  | 61.0837 | 12.4122 | 10 | 59 | 59 | 0.060012 | -0.43999 | 61.0837 |  | 31.8 | 30
|  |  |  | 61.1837 | 12.4325 | 10 | 59 | 59 | 0.058805 | -0.4412 | 61.1837 |  | 31.9 | 30
|  |  |  | 61.2837 | 12.4528 | 10 | 59 | 59 | 0.057622 | -0.44238 | 61.2837 |  | 32 | 30
|  |  |  | 61.3837 | 12.4732 | 10 | 59 | 59 | 0.056463 | -0.44354 | 61.3837 |  | 32.1 | 30
|  |  |  | 61.4837 | 12.4935 | 10 | 59 | 59 | 0.055327 | -0.44467 | 61.4837 |  | 32.2 | 30
|  |  |  | 61.5837 | 12.5138 | 15 | 121 | 121 | 0.11119 | -0.38881 | 61.5837 |  | 32.3 | 30
|  |  |  | 61.6837 | 12.5341 | 15 | 121 | 121 | 0.10895 | -0.39105 | 61.6837 |  | 32.4 | 30
|  |  |  | 61.7837 | 12.5544 | 15 | 121 | 121 | 0.10676 | -0.39324 | 61.7837 |  | 32.5 | 35
|  |  |  | 61.8837 | 12.5748 | 15 | 121 | 121 | 0.10461 | -0.39539 | 61.8837 |  | 32.6 | 35
|  |  |  | 61.9837 | 12.5951 | 15 | 121 | 121 | 0.10251 | -0.39749 | 61.9837 |  | 32.7 | 35
|  |  |  | 62.0837 | 12.6154 | 15 | 121 | 121 | 0.10044 | -0.39956 | 62.0837 |  | 32.8 | 35
|  |  |  | 62.1837 | 12.6357 | 15 | 121 | 121 | 0.098423 | -0.40158 | 62.1837 |  | 32.9 | 35
|  |  |  | 62.2837 | 12.656 | 15 | 121 | 121 | 0.096444 | -0.40356 | 62.2837 |  | 33 | 35
|  |  |  | 62.3837 | 12.6764 | 15 | 121 | 121 | 0.094504 | -0.4055 | 62.3837 |  | 33.1 | 35
|  |  |  | 62.4837 | 12.6967 | 15 | 121 | 121 | 0.092603 | -0.4074 | 62.4837 |  | 33.2 | 35
|  |  |  | 62.5837 | 12.717 | 15 | 121 | 121 | 0.09074 | -0.40926 | 62.5837 |  | 33.3 | 35
|  |  |  | 62.6837 | 12.7373 | 15 | 121 | 121 | 0.088915 | -0.41109 | 62.6837 |  | 33.4 | 35
|  |  |  | 62.7837 | 12.7576 | 15 | 121 | 121 | 0.087126 | -0.41287 | 62.7837 |  | 33.5 | 35
|  |  |  | 62.8837 | 12.778 | 15 | 121 | 121 | 0.085374 | -0.41463 | 62.8837 |  | 33.6 | 35
|  |  |  | 62.9837 | 12.7983 | 15 | 121 | 121 | 0.083656 | -0.41634 | 62.9837 |  | 33.7 | 35
|  |  |  | 63.0837 | 12.8186 | 15 | 121 | 121 | 0.081974 | -0.41803 | 63.0837 |  | 33.8 | 35
|  |  |  | 63.1837 | 12.8389 | 15 | 121 | 121 | 0.080325 | -0.41968 | 63.1837 |  | 33.9 | 35
|  |  |  | 63.2837 | 12.8592 | 15 | 121 | 121 | 0.078709 | -0.42129 | 63.2837 |  | 34 | 35
|  |  |  | 63.3837 | 12.8796 | 15 | 121 | 121 | 0.077126 | -0.42287 | 63.3837 |  | 34.1 | 35
|  |  |  | 63.4837 | 12.8999 | 15 | 121 | 121 | 0.075574 | -0.42443 | 63.4837 |  | 34.2 | 35
|  |  |  | 63.5837 | 12.9202 | 15 | 121 | 121 | 0.074054 | -0.42595 | 63.5837 |  | 34.3 | 35
|  |  |  | 63.6837 | 12.9405 | 15 | 121 | 121 | 0.072565 | -0.42744 | 63.6837 |  | 34.4 | 35
|  |  |  | 63.7837 | 12.9608 | 15 | 121 | 121 | 0.071105 | -0.42889 | 63.7837 |  | 34.5 | 35
|  |  |  | 63.8837 | 12.9812 | 15 | 121 | 121 | 0.069675 | -0.43033 | 63.8837 |  | 34.6 | 35
|  |  |  | 63.9837 | 13.0015 | 15 | 121 | 121 | 0.068273 | -0.43173 | 63.9837 |  | 34.7 | 35
|  |  |  | 64.0837 | 13.0218 | 15 | 121 | 121 | 0.0669 | -0.4331 | 64.0837 |  | 34.8 | 35
|  |  |  | 64.1837 | 13.0421 | 15 | 121 | 121 | 0.065554 | -0.43445 | 64.1837 |  | 34.9 | 35
|  |  |  | 64.2837 | 13.0624 | 15 | 121 | 121 | 0.064236 | -0.43576 | 64.2837 |  | 35 | 35
|  |  |  | 64.3837 | 13.0828 | 15 | 121 | 121 | 0.062944 | -0.43706 | 64.3837 |  | 35.1 | 35
|  |  |  | 64.4837 | 13.1031 | 15 | 121 | 121 | 0.061677 | -0.43832 | 64.4837 |  | 35.2 | 35
|  |  |  | 64.5837 | 13.1234 | 15 | 121 | 121 | 0.060437 | -0.43956 | 64.5837 |  | 35.3 | 35
|  |  |  | 64.6837 | 13.1437 | 15 | 121 | 121 | 0.059221 | -0.44078 | 64.6837 |  | 35.4 | 35
|  |  |  | 64.7837 | 13.164 | 15 | 121 | 121 | 0.05803 | -0.44197 | 64.7837 |  | 35.5 | 35
|  |  |  | 64.8837 | 13.1844 | 15 | 121 | 121 | 0.056863 | -0.44314 | 64.8837 |  | 35.6 | 35
|  |  |  | 64.9837 | 13.2047 | 15 | 121 | 121 | 0.055719 | -0.44428 | 64.9837 |  | 35.7 | 35
|  |  |  | 65.0837 | 13.225 | 15 | 121 | 121 | 0.054598 | -0.4454 | 65.0837 |  | 35.8 | 35
|  |  |  | 65.1837 | 13.2453 | 15 | 121 | 121 | 0.0535 | -0.4465 | 65.1837 |  | 35.9 | 35
|  |  |  | 65.2837 | 13.2656 | 15 | 121 | 121 | 0.052424 | -0.44758 | 65.2837 |  | 36 | 35
|  |  |  | 65.3837 | 13.286 | 15 | 121 | 121 | 0.051369 | -0.44863 | 65.3837 |  | 36.1 | 35
|  |  |  | 65.4837 | 13.3063 | 15 | 121 | 121 | 0.050336 | -0.44966 | 65.4837 |  | 36.2 | 35
|  |  |  | 65.5837 | 13.3266 | 15 | 121 | 121 | 0.049323 | -0.45068 | 65.5837 |  | 36.3 | 35
|  |  |  | 65.6837 | 13.3469 | 15 | 121 | 121 | 0.048331 | -0.45167 | 65.6837 |  | 36.4 | 35
|  |  |  | 65.7837 | 13.3672 | 15 | 121 | 121 | 0.047359 | -0.45264 | 65.7837 |  | 36.5 | 35
|  |  |  | 65.8837 | 13.3876 | 15 | 121 | 121 | 0.046406 | -0.45359 | 65.8837 |  | 36.6 | 35
|  |  |  | 65.9837 | 13.4079 | 15 | 121 | 121 | 0.045473 | -0.45453 | 65.9837 |  | 36.7 | 35
|  |  |  | 66.0837 | 13.4282 | 15 | 121 | 121 | 0.044558 | -0.45544 | 66.0837 |  | 36.8 | 35
|  |  |  | 66.1837 | 13.4485 | 15 | 121 | 121 | 0.043662 | -0.45634 | 66.1837 |  | 36.9 | 35
|  |  |  | 66.2837 | 13.4688 | 15 | 121 | 121 | 0.042784 | -0.45722 | 66.2837 |  | 37 | 35
|  |  |  | 66.3837 | 13.4892 | 15 | 121 | 121 | 0.041923 | -0.45808 | 66.3837 |  | 37.1 | 35
|  |  |  | 66.4837 | 13.5095 | 15 | 121 | 121 | 0.04108 | -0.45892 | 66.4837 |  | 37.2 | 35
|  |  |  | 66.5837 | 13.5298 | 15 | 121 | 121 | 0.040254 | -0.45975 | 66.5837 |  | 37.3 | 35
|  |  |  | 66.6837 | 13.5501 | 15 | 121 | 121 | 0.039444 | -0.46056 | 66.6837 |  | 37.4 | 35
|  |  |  | 66.7837 | 13.5704 | 15 | 121 | 121 | 0.03865 | -0.46135 | 66.7837 |  | 37.5 | 40
|  |  |  | 66.8837 | 13.5908 | 15 | 121 | 121 | 0.037873 | -0.46213 | 66.8837 |  | 37.6 | 40
|  |  |  | 66.9837 | 13.6111 | 15 | 121 | 121 | 0.037111 | -0.46289 | 66.9837 |  | 37.7 | 40
|  |  |  | 67.0837 | 13.6314 | 15 | 121 | 121 | 0.036365 | -0.46364 | 67.0837 |  | 37.8 | 40
|  |  |  | 67.1837 | 13.6517 | 15 | 121 | 121 | 0.035633 | -0.46437 | 67.1837 |  | 37.9 | 40
|  |  |  | 67.2837 | 13.672 | 15 | 121 | 121 | 0.034916 | -0.46508 | 67.2837 |  | 38 | 40
|  |  |  | 67.3837 | 13.6924 | 15 | 121 | 121 | 0.034214 | -0.46579 | 67.3837 |  | 38.1 | 40
|  |  |  | 67.4837 | 13.7127 | 15 | 121 | 121 | 0.033526 | -0.46647 | 67.4837 |  | 38.2 | 40
|  |  |  | 67.5837 | 13.733 | 15 | 121 | 121 | 0.032852 | -0.46715 | 67.5837 |  | 38.3 | 40
|  |  |  | 67.6837 | 13.7533 | 15 | 121 | 121 | 0.032191 | -0.46781 | 67.6837 |  | 38.4 | 40
|  |  |  | 67.7837 | 13.7736 | 15 | 121 | 121 | 0.031543 | -0.46846 | 67.7837 |  | 38.5 | 40
|  |  |  | 67.8837 | 13.794 | 15 | 121 | 121 | 0.030909 | -0.46909 | 67.8837 |  | 38.6 | 40
|  |  |  | 67.9837 | 13.8143 | 15 | 121 | 121 | 0.030287 | -0.46971 | 67.9837 |  | 38.7 | 40
|  |  |  | 68.0837 | 13.8346 | 15 | 121 | 121 | 0.029678 | -0.47032 | 68.0837 |  | 38.8 | 40
|  |  |  | 68.1837 | 13.8549 | 15 | 121 | 121 | 0.029081 | -0.47092 | 68.1837 |  | 38.9 | 40
|  |  |  | 68.2837 | 13.8752 | 15 | 121 | 121 | 0.028496 | -0.4715 | 68.2837 |  | 39 | 40
|  |  |  | 68.3837 | 13.8956 | 15 | 121 | 121 | 0.027923 | -0.47208 | 68.3837 |  | 39.1 | 40
|  |  |  | 68.4837 | 13.9159 | 15 | 121 | 121 | 0.027361 | -0.47264 | 68.4837 |  | 39.2 | 40
|  |  |  | 68.5837 | 13.9362 | 15 | 121 | 121 | 0.026811 | -0.47319 | 68.5837 |  | 39.3 | 40
|  |  |  | 68.6837 | 13.9565 | 15 | 121 | 121 | 0.026271 | -0.47373 | 68.6837 |  | 39.4 | 40
|  |  |  | 68.7837 | 13.9768 | 15 | 121 | 121 | 0.025743 | -0.47426 | 68.7837 |  | 39.5 | 40
|  |  |  | 68.8837 | 13.9972 | 15 | 121 | 121 | 0.025225 | -0.47477 | 68.8837 |  | 39.6 | 40
|  |  |  | 68.9837 | 14.0175 | 15 | 121 | 121 | 0.024718 | -0.47528 | 68.9837 |  | 39.7 | 40
|  |  |  | 69.0837 | 14.0378 | 15 | 121 | 121 | 0.02422 | -0.47578 | 69.0837 |  | 39.8 | 40
|  |  |  | 69.1837 | 14.0581 | 15 | 121 | 121 | 0.023733 | -0.47627 | 69.1837 |  | 39.9 | 40
|  |  |  | 69.2837 | 14.0784 | 15 | 121 | 121 | 0.023256 | -0.47674 | 69.2837 |  | 40 | 40
|  |  |  | 69.3837 | 14.0988 | 15 | 121 | 121 | 0.022788 | -0.47721 | 69.3837 |  | 40.1 | 40
|  |  |  | 69.4837 | 14.1191 | 15 | 121 | 121 | 0.02233 | -0.47767 | 69.4837 |  | 40.2 | 40
|  |  |  | 69.5837 | 14.1394 | 15 | 121 | 121 | 0.021881 | -0.47812 | 69.5837 |  | 40.3 | 40
|  |  |  | 69.6837 | 14.1597 | 15 | 121 | 121 | 0.02144 | -0.47856 | 69.6837 |  | 40.4 | 40
|  |  |  | 69.7837 | 14.18 | 15 | 121 | 121 | 0.021009 | -0.47899 | 69.7837 |  | 40.5 | 40
|  |  |  | 69.8837 | 14.2004 | 15 | 121 | 121 | 0.020587 | -0.47941 | 69.8837 |  | 40.6 | 40
|  |  |  | 69.9837 | 14.2207 | 15 | 121 | 121 | 0.020172 | -0.47983 | 69.9837 |  | 40.7 | 40
|  |  |  | 70.0837 | 14.241 | 15 | 121 | 121 | 0.019767 | -0.48023 | 70.0837 |  | 40.8 | 40
|  |  |  | 70.1837 | 14.2613 | 15 | 121 | 121 | 0.019369 | -0.48063 | 70.1837 |  | 40.9 | 40
|  |  |  | 70.2837 | 14.2816 | 15 | 121 | 121 | 0.018979 | -0.48102 | 70.2837 |  | 41 | 40
|  |  |  | 70.3837 | 14.302 | 15 | 121 | 121 | 0.018598 | -0.4814 | 70.3837 |  | 41.1 | 40
|  |  |  | 70.4837 | 14.3223 | 15 | 121 | 121 | 0.018224 | -0.48178 | 70.4837 |  | 41.2 | 40
|  |  |  | 70.5837 | 14.3426 | 15 | 121 | 121 | 0.017857 | -0.48214 | 70.5837 |  | 41.3 | 40
|  |  |  | 70.6837 | 14.3629 | 15 | 121 | 121 | 0.017498 | -0.4825 | 70.6837 |  | 41.4 | 40
|  |  |  | 70.7837 | 14.3832 | 15 | 121 | 121 | 0.017146 | -0.48285 | 70.7837 |  | 41.5 | 40
|  |  |  | 70.8837 | 14.4036 | 15 | 121 | 121 | 0.016801 | -0.4832 | 70.8837 |  | 41.6 | 40
|  |  |  | 70.9837 | 14.4239 | 15 | 121 | 121 | 0.016463 | -0.48354 | 70.9837 |  | 41.7 | 40
|  |  |  | 71.0837 | 14.4442 | 15 | 121 | 121 | 0.016132 | -0.48387 | 71.0837 |  | 41.8 | 40
|  |  |  | 71.1837 | 14.4645 | 15 | 121 | 121 | 0.015807 | -0.48419 | 71.1837 |  | 41.9 | 40
|  |  |  | 71.2837 | 14.4848 | 15 | 121 | 121 | 0.015489 | -0.48451 | 71.2837 |  | 42 | 40
|  |  |  | 71.3837 | 14.5052 | 15 | 121 | 121 | 0.015178 | -0.48482 | 71.3837 |  | 42.1 | 40
|  |  |  | 71.4837 | 14.5255 | 15 | 121 | 121 | 0.014873 | -0.48513 | 71.4837 |  | 42.2 | 40
|  |  |  | 71.5837 | 14.5458 | 15 | 121 | 121 | 0.014573 | -0.48543 | 71.5837 |  | 42.3 | 40
|  |  |  | 71.6837 | 14.5661 | 15 | 121 | 121 | 0.01428 | -0.48572 | 71.6837 |  | 42.4 | 40

## Sheet7 (hidden)
Final Results
radioisotope: | Fluorine -18
|  | combined code = |  | 2221 | calculation type code= |  | 20
shielding material : | concrete |  |  |  | build up factor code = |  | 1
geometry: | 4π | 1111 | --
|  | 1112 | --
calculation type: | thickness of shielding required | 1211 | --
|  | 1212 | --
calculation with /without build up factor : |  | 2111 | --
| with | 2112 | -- | final exposure W or W/O build up factor = |  | ---
|  | 2211 | --
|  | 2212 | -- | final thickness W or W/O build up factor required = |  |  | 50.6837
|  | 3111 | --
|  | 3112 | -- | final result -final exposure or final thickness W or W/O build up factor= |  |  |  | 50.6837
|  | 3211 | --
|  | 3212 | --
|  | 1121 | ----
|  | 1122 | ----
|  | 1221 | ----
|  | 1222 | ----
|  | 2121 | ----
|  | 2122 | ----
|  | 2221 | 50.6837
|  | 2222 | ----
|  | 3121 | ----
|  | 3122 | ----
|  | 3221 | ----
|  | 3222 | ----

## Reference
References:
Oak Ridge National Laboratory publication 45 (ORNL / RSIC -45 /R1)
National Institute of Standards and Technology (NIST) :
physical reference data X-ray and gamma ray data table 3
American Nuclear Society : Takeuchi K, Tanaka S-I, Nucl Sci Eng, 87 :4, 478-489
Radiological Health Handbook, US Department of Health, pp 145-149

## Geometry
Geometry :
In radiation shielding, geometry describes the position of the detector relative to the radioactive source.
2 π geometry refers to situations where the shielding intercepts radiation emitted from only one side of the source (e.g., a wall barrier).
4 π geometry refers to situations where the shielding intercepts radiation emitted from all directions (e.g., a vial shield).

## Build up factor
Build up factor (B) :
The basic shielding equation assumes that the scattered photons will not re-enter the field of the radiation detector and hence will not contribute to the radiation exposure measured
In reality, many scattered photons do re-enter the field of the detector and contribute significantly to the exposure
Build up factor is a unit-less factor that accounts for scattered photons re-entering the field
This calculator applies build up factor when selected
The value of build up factor depends on
Density of shielding material – the higher the density , the larger the build up factor
Photon energy – the higher the photon energy , the larger the build up factor
Thickness of the shielding – the thicker the shielding, the larger the build up factor
One would not know the build up factor value unless one already knows the thickness of the shielding . However, such thickness is unknow and the objective of the shielding equation is to determine the thickness required
Determination of build up factor
By experiment - by measuring the exposure from photons of different energies through shielding of different density at different thickness
By simulation - it is possible to determine the build up factors by different simulation techniques such as Monte Carlo simulations, MCNP, etc

## About Sidney Yu
Sidney Yu, PhD, FRSC
Sidney Yu is a semi-retired radiopharmacist with over 30 years of experience in nuclear medicine and radiopharmacy. He obtained his PhD from Purdue University, College of Pharmacy and Pharmaceutical Sciences, and also holds a certificate in Radiation Health Physics and Radiation Engineering from The Ohio State University.
In 2025, He was elected a Fellow of the Royal Society of Chemistry (FRSC) , in recognition of his contributions to the chemical sciences and professional leadership.
He began his career as a radiopharmacist and research scientist at the Singapore General Hospital, Department of Nuclear Medicine and PET, where he worked from 1995 to 2018. During this time, he established a centralised radiopharmacy service supporting multiple hospitals, led the establishment of the cyclotron facility and GMP approval process, and expanded radiopharmacy services to support Radioembolisation , Peptide Receptor Radionuclide Therapy, and other radioisotope therapies. His work covered all aspects of radiopharmacy operations, clinical service development, radiation safety, teaching, and research.
Since 2018, Sidney has served as the Director of Radiopharmaceuticals with Singapore Advanced Medicine Imaging, where he oversees radiopharmacy operations and radiation safety for the facility. He is actively involved in radiation protection, shielding design, regulatory compliance, and exposure calculations
Contact
meditopes@yahoo.com